PAGE  
August 2013
DQ-DR Combi Tray
Page 1 of 9
Rev. No: 01
57R
Worksheet can be downloaded from


Worksheet
www.olerup-ssp.com










	PHOTO DOCUMENT


Olerup SSP®
Olerup SSP® DQ-DR Combi Tray
Prod. No:
101.704-48/48u

101.704-12/12u
Lot No:    
57R
Expiry Date:    
2015-July-01
Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 8 primer mixes used for DQ low resolution SSP typing.
	Primer Mix
	Size of spec. PCR product1
	Size of control band2
	DQ serology3
	Amplified DQB1 alleles4

	1
	225 bp
	515 bp
	5
	*05:01:01:01-05:18

	2
	220 bp, 270 bp
	430 bp
	1, 5, 6, null
	*06:01:01-06:52

	3
	210 bp
	430 bp
	2
	*02:01:01-02:07

	46
	220 bp
	515 bp
	3, 7
	*03:01:01:01-03:01:06, 03:04, 03:09-03:10, 03:13-03:14, 03:16, 03:19, 03:21-03:22, 03:24, 03:27-03:29, 03:35-03:36, 03:42, 03:44

	56
	130 bp
	515 bp
	6, 8
	*03:02:01-03:02:05, 03:05:01-03:05:04, 03:07-03:08, 03:11, 03:18, 03:32, 03:37, 06:29

	66,7
	135 bp
	515 bp
	2, 3, 4, 9
	*02:03, 03:03:02:01-03:03:04, 03:06, 03:12, 03:15, 03:20, 03:25-03:26, 03:30-03:31, 03:33-03:34, 03:38-03:41, 03:43, 04:03:01-04:03:02, 06:51

	75,6
	145 bp, 185 bp
	515 bp
	3, 7, 8, 9
	*03:01:01:01-03:44

	86
	160 bp
	430 bp
	4
	*04:01:01-04:08


1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DQ low resolution SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given. 

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DQ low resolution typing.

In addition, wells number 4, 5, 6 and 7 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

 In the presence of a specific amplification the intensity of the control band often decreases.

3The serological reactivity of all DQ alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele. The DQB1*03:10 allele has been assigned type DQ7 by NMDP.
4For several DQB1 alleles 1st and/or 3rd exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups. 
5Primer mix 7 may have tendencies of unspecific amplifications.

6Primer mixes 4 to 8 may give rise to a lower yield of HLA-specific PCR product than the other DQ low primer mixes.

7Primer mix 6 may have a tendency to giving rise to primer oligomer formation.
‘w’, may be weakly amplified.

[image: image1.emf]Well No. 1 2 3 4 5 6 7 8

DQB1 allele

1

ser.

2

*05:01:01:01-05:18 DQ5, - 1

*06:01:01-06:28, 06:30-06:50, 06:52

DQ1, DQ5, 

DQ6, Null, -

2

*06:29 DQ6 2 5

*06:51 - 2 6

*02:01:01-02:02, 02:04-02:07 DQ2, - 3

*02:03 DQ2 3 6

*03:01:01:01-03:01:06, 03:04, 03:09-03:10, 03:13-03:14, 

03:16, 03:19, 03:21-03:22, 03:24, 03:27-03:29, 03:35-03:36, 

03:42, 03:44

DQ3, DQ7, - 4 7

*03:02:01-03:02:05, 03:05:01-03:05:04, 03:07-03:08, 03:11, 

03:18, 03:32, 03:37

DQ8, - 5 7

*03:03:02:01-03:03:04, 03:06, 03:12, 03:15, 03:20, 03:25-

03:26, 03:30-03:31, 03:33-03:34, 03:38-03:41, 03:43

DQ3,DQ9,  - 6 7

*03:17, 03:23 DQ9,  - 7

*04:01:01-04:02:02, 04:04-04:08 DQ4, - 8

*04:03:01-04:03:02 - 6 8

Length of spec. PCR product(s)

225220210220130135145160

270 185

Well No. 1 2 3 4 5 6 7 8

Specific Reaction

1st match: DQB1*

2nd match: DQB1*

Control Reaction

Failed Controls

False Positive

False Negative


1The sequence of the DQB1*030301 allele has been shown to be identical to DQB1*03:03:02.

2The serological reactivity of all DQ alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele. The DQB1*03:10 allele has been assigned type DQ7 by NMDP.
 ‘w’, might be weakly amplified.

Specificities and sizes of the PCR products of the 23+1 primer mixes used for DR low resolution SSP typing.
	Primer Mix
	Size of spec. PCR product1
	Size of control band2
	DR

serology3
	Amplified HLA-DRB alleles4

	  96,8
	210 bp, 

260 bp
	515 bp
	1
	*01:01:01-01:02:08, 01:04-01:21, 01:22w, 01:23-01:38, 01:40N-01:53

	10
	200 bp
	430 bp
	103
	*01:03, 01:39N, 01:42

	116
	200 bp, 

220 bp
	430 bp
	2, 15
	*15:01:01:01-15:81

	12
	210 bp
	430 bp
	16
	*16:01:01-16:05:02, 16:07-16:19

	135,6,7
	120 bp, 

220 bp
	430 bp
	3, 17, 18, 11
	*03:01:01:01-03:75, 03:77-03:84, 03:86, 11:07, 11:53, 11:103, 11:105, 11:107, 11:125, 15:25

	145,6,7
	80 bp, 

210 bp
	430 bp
	3, 6, 17, 

11, 13, 14
	*03:01:01:01-03:01:20, 03:04:01-03:06, 03:08-03:16, 03:18-03:20, 03:22-03:23, 03:25-03:26, 03:28, 03:30-03:31, 03:33-03:34, 03:36-03:37, 03:43-03:48, 03:50-03:52, 03:54-03:68N, 03:70-03:73, 03:75-03:84, 03:86, 08:40, 11:02:01-11:03, 11:11:01-11:11:02, 11:14:01-11:14:02, 11:16, 11:20-11:21, 11:36, 11:40-11:41, 11:48, 11:59, 11:63, 11:65:01-11:65:02, 11:68, 11:70, 11:73, 11:76, 11:79-11:80, 11:83, 11:85-11:87, 11:93, 11:118, 11:122, 11:124, 11:127, 11:131-11:132, 13:01:01-13:04, 13:08, 13:10, 13:15-13:17, 13:19-13:20, 13:22-13:24, 13:27-13:29, 13:31-13:41, 13:43, 13:45, 13:48, 13:51-13:54, 13:57, 13:59, 13:61:01-13:61:02, 13:63-13:66:02, 13:68-13:76, 13:78-13:81, 13:83-13:85, 13:87-13:99, 13:101-13:102, 13:104-13:107, 13:109, 13:111-13:117, 13:120-13:131, 13:133, 13:135, 13:137N-13:143, 14:16, 14:19, 14:21, 14:82, 14:95, 14:109

	155,6
	85 bp, 

210 bp
	430 bp
	3, 6, 11, 13, 14, 1403, 18
	*03:02:01-03:03, 03:27, 03:29, 03:38, 03:53, 03:74, 11:13:01w-11:13:02w, 11:26, 11:34, 13:15, 13:19, 13:26:01-13:26:02, 13:44, 13:53, 13:57, 13:85-13:86, 13:104, 14:02-14:03:02, 14:06:01-14:06:02, 14:09, 14:12:01-14:13, 14:17-14:21, 14:24, 14:27, 14:29-14:30, 14:32:01w-14:32:02w, 14:33, 14:40-14:41, 14:47-14:49, 14:51, 14:63, 14:65w, 14:67, 14:77-14:78, 14:80-14:81, 14:83, 14:85, 14:89, 14:94, 14:98, 14:102, 14:106, 14:108-14:109, 14:115, 14:119, 14:121

	165,6,8
	100 bp, 

175 bp
	430 bp
	4
	*04:01:01-04:05:11, 04:05:13-04:133

	  
	
	
	
	

	176
	210 bp, 

235 bp
	430 bp
	7, 13, 14
	*07:01:01:01-07:01:04, 07:03-07:22, 12:22, 13:17, 13:116, 14:50

	186
	170 bp, 215 bp, 250 bp
	515 bp
	8, 12, 14
	*08:01:01-08:19, 08:21-08:49, 11:67, 12:04, 12:16:01, 12:22, 14:11, 14:15, 14:68, 14:93

	195,6
	90 bp, 135 bp, 180 bp 
	430 bp
	3, 9, 11
	*03:08, 03:65, 09:01:02-09:20, 11:07, 11:53, 11:103, 11:105, 11:107, 11:125

	20
	175 bp
	430 bp
	10, 11, 13
	*03:76, 10:01:01-10:04, 11:59, 11:80, 11:83, 11:87, 13:27, 13:41, 13:71, 13:129

	215,6
	100 bp, 

170 bp
	430 bp
	3, 8, 11, 14
	*03:08, 03:65, 08:31, 08:41, 11:01:01-11:70, 11:72-11:129, 11:131-11:134

	225,6
	85 bp, 

105 bp
	430 bp
	12
	*08:32, 12:01:01-12:37

	238
	215 bp
	430 bp
	6, 8, 11, 13, 14, 1403
	*03:76, 08:20-08:21, 11:01:01-11:04:10, 11:06:01-11:06:03, 11:08:01-11:12:02, 11:14:01-11:16, 11:18-11:21, 11:23-11:25, 11:27:01-11:33, 11:35-11:51, 11:54:01-11:54:02, 11:56-11:66, 11:68, 11:70, 11:72-11:81, 11:83-11:88, 11:90-11:97, 11:99-11:102:02, 11:106, 11:108-11:124, 11:126-11:129, 11:131-11:134, 13:01:01-13:02:01, 13:02:03-13:08, 13:10-13:16, 13:18-13:43, 13:45-13:85, 13:87-13:115, 13:117-13:128, 13:130-13:143, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:16, 14:19, 14:21-14:22, 14:25, 14:27, 14:40, 14:53, 14:63, 14:67, 14:69, 14:74, 14:77-14:78, 14:84-14:85, 14:98, 14:102, 14:105, 14:109, 14:115-14:116, 14:128, DRB3*02:27

	246,8
	195 bp, 

225 bp
	430 bp
	6, 8, 11, 12, 13, 14
	*08:01:01-08:02:04, 08:04:01-08:09, 08:11, 08:16-08:17, 08:20-08:22, 08:24, 08:26, 08:28, 08:31, 08:39, 08:41-08:44, 11:01:01-11:06:03, 11:09-11:12:02, 11:14:01-11:16, 11:20-11:21, 11:23-11:25, 11:27:01-11:30, 11:32-11:33, 11:35-11:41, 11:43-11:44, 11:46:01-11:51, 11:54:01-11:56, 11:58:01-11:63, 11:65:01-11:70, 11:72, 11:74:01-11:78, 11:80-11:88, 11:90-11:97, 11:99-11:102:02, 11:106, 11:108-11:118, 11:120-11:124, 11:126-11:129, 11:133-11:134, 12:02:01-12:02:05, 12:13, 12:15-12:16:01, 12:18-12:21, 12:23, 12:26-12:27, 12:31N-12:33, 12:37, 13:01:01-13:02:01, 13:02:03-13:02:07, 13:04-13:05:02, 13:07:01-13:09, 13:11:01-13:11:02, 13:14:01-13:24, 13:26:01-13:29, 13:31-13:32, 13:34-13:36, 13:38-13:43, 13:45-13:55, 13:57, 13:59, 13:61:01-13:65, 13:67-13:76, 13:78-13:80, 13:83-13:84, 13:87, 13:91-13:93, 13:96:01-13:100, 13:102-13:109, 13:111-13:114, 13:116-13:117, 13:121, 13:123-13:132, 13:135-13:136, 13:138-13:143, 14:15-14:16, 14:22, 14:24-14:25, 14:27, 14:37, 14:53, 14:73, 14:105, 14:128

	
	
	
	
	

	257
	175 bp
	430 bp
	3, 6, 11, 13, 14, 1403, 17, 18
	*03:01:01:01-03:07, 03:09, 03:11:01-03:41, 03:43-03:45, 03:47-03:63, 03:66-03:68N, 03:70-03:84, 03:86, 08:20, 13:01:01-13:16, 13:18-13:42, 13:44, 13:46-13:66:02, 13:68-13:102, 13:104-13:115, 13:117-13:121, 13:123-13:143, 14:02-14:03:02, 14:05:01-14:06:02, 14:09, 14:12:01-14:14, 14:17-14:21, 14:23:01, 14:23:03-14:24, 14:27, 14:29-14:30, 14:33, 14:36-14:37, 14:40-14:45, 14:47-14:48, 14:51, 14:56, 14:59, 14:63-14:65, 14:67, 14:77-14:78, 14:80-14:81, 14:83-14:85, 14:89, 14:91, 14:94-14:96, 14:98, 14:100, 14:102-14:103, 14:106, 14:108-14:109, 14:115-14:116, 14:121, 14:123, 14:127

	265,6
	100 bp, 140 bp, 155 bp 
	430 bp
	4, 6, 8,13, 14, 1404
	*04:62, 04:69, 04:73, 04:105:01-04:105:02, 04:122, 08:08, 11:69, 11:82, 13:45, 14:01:01-14:01:02, 14:04, 14:07:01-14:07:02, 14:10, 14:16, 14:22, 14:25-14:26, 14:28, 14:31-14:32:02, 14:35, 14:37-14:39, 14:49-14:50, 14:53-14:54:01, 14:55, 14:57-14:58, 14:60-14:62, 14:68-14:71, 14:73-14:76, 14:79, 14:82, 14:86-14:88, 14:90, 14:93, 14:99, 14:101, 14:104-14:105, 14:107, 14:110-14:114, 14:117-14:120, 14:122, 14:124-14:125, 14:128-14:129, 14:131, DRB4*01:03:01:02N

	275,6,9
	110 bp, 135 bp, 170 bp
	430 bp
	3, 4, 6, 9, 11, 13,14,1404


	*03:10, 09:01:02-09:01:05, 09:01:07-09:02:02, 09:04-09:20, 11:13:01-11:13:02, 11:17, 11:52, 13:43, 14:01:01-14:02, 14:04-14:11, 14:13-14:14, 14:16-14:18, 14:19w, 14:20, 14:21w, 14:22-14:23:03, 14:26, 14:28-14:36, 14:38-14:39, 14:41, 14:43-14:52, 14:54:01-14:57, 14:59-14:62, 14:64-14:65, 14:68, 14:70-14:76, 14:79-14:83, 14:86-14:88, 14:90-14:97, 14:99-14:101, 14:103-14:108, 14:109w, 14:110-14:114, 14:117-14:127, 14:129-14:131, 15:27, 15:34, 15:66

	285,6,8
	110 bp, 175 bp, 225 bp
	430 bp
	2, 3, 4, 6, 8, 11,13, 14, 1403, 1404, 16


	*03:10, 08:09, 08:20-08:21, 08:32, 08:35, 08:36:02, 11:13:01-11:13:02, 11:17, 11:23, 11:25, 11:31, 11:45, 11:52, 11:55, 11:64, 11:89, 11:96, 11:119, 13:13, 13:18, 13:43, 13:45, 13:47, 13:55, 13:119, 14:01:01-14:01:03, 14:03:01-14:05:04, 14:07:01-14:08, 14:10-14:12:02, 14:14-14:16, 14:18, 14:22-14:23:03, 14:25-14:28, 14:31-14:32:02, 14:34-14:36, 14:38-14:40, 14:42-14:45, 14:49-14:50, 14:53-14:65, 14:67-14:79, 14:81-14:82, 14:84-14:93, 14:95-14:97, 14:99-14:105, 14:107, 14:110-14:120, 14:122-14:131, 15:21w, 16:04w, 16:18w

	296,7
	160 bp, 

240 bp
	430 bp
	52
	DRB3*01:01:02:01-01:15, DRB3*02:01-02:29N, DRB3*03:01:01-03:03

	308,10
	215 bp
	430 bp
	53
	DRB4*01:01:01:01-01:08

	31
	175 bp
	430 bp
	51
	DRB5*01:01:01-01:14, DRB5*02:02-02:06

	3211
	-
	-
	
	Negative Control


1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DR low resolution SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given. 

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers, e.g. the primers in wells 11, 26, 27 and 28.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 9 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DR low resolution typing.

In addition, well number 18 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

 In the presence of a specific amplification the intensity of the control band often decreases.

3The serological split of all DRB1 alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele. 

4For several DRB1 alleles 1st and/or 3rd exon(s) and beyond, as well as intron nucleotide sequences, are not available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in these regions are conserved within allelic groups. 
5Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

6Individual alleles can give to rise to two differently sized specific PCR fragments in primer mixes 9, 11, 13 to 19, 21, 22, 24 and 26 to 29.

7Due to sharing of sequence motifs in codon 38, DRB3*01:14 will also be amplified in primer mixes 13, 14 and 25 in addition to primer mix 29.

8Primer mixes 9, 16, 23, 24, 28 and 30 may have a tendency to giving rise to primer oligomer formation.

9Primer mix 27 has a tendency of giving rise to primer oligomer formation and also has an intense primer cloud due to the high number of primers present in the primer mix.
10The DRB4*01:03:01:02N allele is amplified by primer mix 30, whereas the DRB4*02:01N and DRB4*03:01N null alleles are not amplified by this primer pair.

11Primer mix 32 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

‘w’, might be weakly amplified.

[image: image2.emf]Well No. DR 9 1011121314151617181920212223242526272829303132

DRB1 allele

1,2

ser

3

*01:01:01-01:02:08, 01:04-01:21, 01:23-01:38, 

01:40N-01:41, 01:43-01:53

DR1, 

Null, -

9

*01:03, 01:39N

DR1, 

DR103, 

Null

10

*01:22 - w

*01:42 - 9 10

*03:01:01:01-03:01:20, 03:04:01-03:06, 03:09, 

03:11:01-03:16, 03:18-03:20, 03:22-03:23, 

03:25-03:26, 03:28, 03:30-03:31, 03:33-03:34, 

03:36-03:37, 03:43-03:45, 03:47-03:48, 03:50-

03:52, 03:54-03:63, 03:66-03:68N, 03:70-03:73, 

03:75, 03:77-03:84, 03:86

DR17, 

DR3, 

Null, -

1314 25

*03:02:01-03:03, 03:27, 03:29, 03:38, 03:53, 

03:74

DR18, 

DR3, -

13 15 25

*03:07, 03:17, 03:21, 03:24, 03:32, 03:35, 03:39-

03:41, 03:49

DR3, - 13 25

*03:08, 03:65 DR3, - 1314 19 21

*03:10 DR3 1314 2728

*03:42, 03:69 - 13

*03:46, 03:64 - 1314

*03:76 - 14 20 23 25

*04:01:01-04:05:11, 04:05:13-04:61, 04:63-

04:68, 04:70-04:72:02, 04:74-04:104, 04:106-

04:121, 04:123-04:133

DR4, 

DR4/DR

3, Null, -

16

*04:62, 04:69, 04:73, 04:105:01-04:105:02, 

04:122

DR4, - 16 26

*07:01:01:01-07:01:04, 07:03-07:22

DR7, 

Null, -

17

*08:01:01-08:02:04, 08:04:01-08:07, 08:11, 

08:16-08:17, 08:22, 08:24, 08:26, 08:28, 08:39, 

08:42-08:44

DR8, - 18 24

*08:03:02-08:03:03, 08:10, 08:12-08:15, 08:18-

08:19, 08:23, 08:25, 08:27, 08:29-08:30:03, 

08:33-08:34, 08:36:01, 08:37-08:38, 08:45-

08:49

DR8, - 18

*08:08 DR8 18 24 26

*08:09, 14:15 DR8 18 24 28

*08:20, 13:18, 13:47, 13:55 DR13, - 232425 28

*08:21 DR8 18 2324 28

*08:31, 08:41, 11:67

DR8, 

DR11, -

18 21 24

*08:32 - 18 22 28

*08:35, 08:36:02 - 18 28

*08:40 - 14 18

*09:01:02-09:01:05, 09:01:07-09:02:02, 09:04-

09:20

DR9, - 19 27

*09:01:06, 09:03 DR9 19

*10:01:01-10:04 DR10, - 20

*11:01:01-11:01:17, 11:04:01-11:04:10, 

11:06:01-11:06:03, 11:09-11:10:02, 11:12:01-

11:12:02, 11:15, 11:24, 11:27:01-11:30, 11:32-

11:33, 11:35, 11:37:01-11:39, 11:43-11:44, 

11:46:01-11:47, 11:49:01-11:51, 11:54:01-

11:54:02, 11:56, 11:58:01-11:58:02, 11:60-

11:62, 11:66, 11:72, 11:74:01-11:75, 11:77-

11:78, 11:81, 11:84:01-11:84:02, 11:88, 11:90-

11:92, 11:94-11:95, 11:97, 11:99-11:102:02, 

11:106, 11:108-11:117, 11:120-11:121, 11:123, 

11:126, 11:128-11:129, 11:133-11:134

DR11, - 21 2324

Well No. DR 9 1011121314151617181920212223242526272829303132

Negative control



[image: image3.emf]Well No. DR 9 1011121314151617181920212223242526272829303132

*11:02:01-11:03, 11:11:01-11:11:02, 11:14:01-

11:14:02, 11:16, 11:20-11:21, 11:36, 11:40-

11:41, 11:48, 11:63, 11:65:01-11:65:02, 11:68, 

11:70, 11:76, 11:85-11:86, 11:93, 11:118, 

11:122, 11:124, 11:127

DR11, 

DR11/D

R13, -

14 21 2324

*11:05 DR11 21 24

*11:07, 11:53, 11:103, 11:105, 11:107, 11:125 DR11, - 13 19 21

*11:08:01-11:08:02, 11:18-11:19:03, 11:42, 

11:57

DR11, - 21 23

*11:13:01-11:13:02 DR11 w 21 2728

*11:17, 11:52

DR11, 

DR14

21 2728

*11:22, 11:98, 11:104 - 21

*11:23, 11:25, 11:96 DR11, - 21 2324 28

*11:26, 11:34 DR11 15 21

*11:31, 11:45, 11:64, 11:119 DR11, - 21 23 28

*11:55 DR11 21 24 28

*11:59, 11:80, 11:83, 11:87 DR11, - 14 2021 2324

*11:69, 11:82 - 21 24 26

*11:73, 11:79, 11:131-11:132 - 14 21 23

*11:89 - 21 28

*12:01:01-12:01:04, 12:03:02-12:03:03, 12:05-

12:12, 12:14, 12:17, 12:24N-12:25, 12:28-

12:30, 12:34-12:36

DR12, 

Null, -

22

*12:02:01-12:02:05, 12:13, 12:15, 12:18-12:21, 

12:23, 12:26-12:27, 12:31N-12:33, 12:37

DR12, 

Null, -

22 24

*12:04 - 18 22

*12:16:01 - 18 22 24

*12:22 - 1718 22

*13:01:01-13:02:01, 13:02:03-13:02:07, 13:04, 

13:08, 13:16, 13:20, 13:22-13:24, 13:28-13:29, 

13:31-13:32, 13:34-13:36, 13:38-13:40, 13:48, 

13:51-13:52, 13:54, 13:59, 13:61:01-13:61:02, 

13:63-13:65, 13:68-13:70, 13:72-13:76, 13:78-

13:80, 13:83-13:84, 13:87, 13:91-13:93, 

13:96:01-13:99, 13:102, 13:105-13:107, 

13:109, 13:111-13:114, 13:117, 13:121, 13:123-

13:128, 13:130-13:131, 13:135, 13:138-13:143

DR11, 

DR13, 

DR14, 

Null, -

14 232425

*13:02:02 DR13 14 25

*13:03:01-13:03:06, 13:10, 13:33:01-13:33:03, 

13:37, 13:66:01-13:66:02, 13:81, 13:88-13:90, 

13:94-13:95, 13:101, 13:115, 13:120, 13:133, 

13:137N

DR13, 

Null, -

14 23 25

*13:05:01-13:05:02, 13:07:01-13:07:02, 

13:11:01-13:11:02, 13:14:01-13:14:03, 

13:21:01-13:21:02, 13:42, 13:46, 13:49-

13:50:02, 13:62, 13:100, 13:108, 13:132, 

DR13, 

DR13/D

R11, 

DR6, -

232425

*13:06, 13:12:01-13:12:03, 13:25, 13:30, 13:56, 

13:58, 13:60, 13:77, 13:82, 13:110, 13:118, 

13:134

DR11, 

DR13, 

DR6,  -

23 25

*13:09 DR13 2425

*13:13, 13:119, 14:84, 14:116 DR13, - 23 25 28

*13:15, 13:19, 13:53, 13:57, 13:104 DR13, - 1415 232425

*13:17, 13:116 DR13, - 14 17 24

*13:26:01-13:26:02 DR14, - 15 232425

*13:27, 13:41, 13:71 DR13 14 20 232425

*13:43 DR13 14 2324 2728

*13:44, 13:86 - 15 25

*13:45 DR13 14 2324 26 28

Well No. DR 9 1011121314151617181920212223242526272829303132

Negative control



[image: image4.emf]Well No. DR 9 1011121314151617181920212223242526272829303132

*13:67, 13:103 DR13, - 2324

*13:85 - 1415 23 25

*13:122 - 14 23

*13:129 - 14 20 2425

*14:01:01-14:01:02, 14:04, 14:07:01-14:07:02, 

14:10, 14:26, 14:28, 14:31, 14:35, 14:38-14:39, 

14:54:01, 14:55, 14:57, 14:60-14:62, 14:70-

14:71, 14:75-14:76, 14:79, 14:86-14:88, 14:90, 

14:99, 14:101, 14:104, 14:107, 14:110-14:114, 

14:117-14:118, 14:120, 14:122, 14:124-14:125, 

14:129, 14:131

DR14, 

DR1404

/DR14,D

R4/DR6,  

-

262728

*14:01:03, 14:08, 14:23:02, 14:34, 14:54:02, 

14:72, 14:92N, 14:97, 14:126, 14:130

DR14, 

Null, -

2728

*14:02, 14:06:01-14:06:02, 14:09, 14:13, 14:17, 

14:20, 14:29-14:30, 14:33, 14:41, 14:47-14:48, 

14:51, 14:80, 14:83, 14:94, 14:106, 14:108, 

14:121

DR14, 

DR6, -

15 25 27

*14:03:01-14:03:02, 14:12:01-14:12:02, 14:40, 

14:63, 14:67, 14:77-14:78, 14:85, 14:102, 

14:115

DR1403

/DR14, 

DR6, -

15 23 25 28

*14:05:01-14:05:04, 14:14, 14:23:01, 14:23:03, 

14:36, 14:43-14:45, 14:56, 14:59, 14:64, 14:91, 

14:96, 14:100, 14:103, 14:123, 14:127

DR14, - 25 2728

*14:11 DR14 18 2728

*14:16 DR6 14 2324 262728

*14:18, 14:81 DR14, - 15 25 2728

*14:19, 14:21, 14:109 DR14, - 1415 23 25 w

*14:22, 14:105 DR14, - 2324 262728

*14:24 DR14 15 2425

*14:25, 14:53, 14:128

DR6,DR

13/14, -

2324 26 28

*14:27 DR14 15 232425 28

*14:32:01-14:32:02 DR14 w 262728

*14:37 DR14 242526

*14:42 - 25 28

*14:46, 14:52 DR14 27

*14:49, 14:119 DR14, - 15 262728

*14:50 DR14 17 262728

*14:58 DR14 26 28

*14:65 DR6 w 25 2728

*14:68, 14:93 DR14, - 18 262728

*14:69 - 23 26 28

*14:73 - 24 262728

*14:74 - 23 262728

*14:82 - 14 262728

*14:89 - 15 25 28

*14:95 - 14 25 2728

*14:98 - 15 23 25

*15:01:01:01-15:20, 15:22-15:24, 15:26, 15:28-

15:33, 15:35-15:65, 15:67-15:81

DR15, 

DR2, 

Null, -

11

*15:21 DR2 11 w

*15:25 - 11 13

*15:27, 15:34, 15:66 - 11 27

*16:01:01-16:03, 16:05:01-16:05:02, 16:07-

16:17, 16:19

DR16, 

Null, -

12

*16:04, 16:18 DR16, - 12 w

Well No. DR 9 1011121314151617181920212223242526272829303132

Negative control



[image: image5.emf]Well No. DR 9 1011121314151617181920212223242526272829303132

DRB3*01:01:02:01-01:15, DRB3*02:01-02:26, 

DRB3*02:28-02:29N, DRB3*03:01:01-03:03

DR52, 

Null, -

29

DRB3*02:27 - 23 29

DRB4*01:01:01:01-01:03:01:01, 

DRB4*01:03:01:03-01:08

DR53, - 30

DRB4*01:03:01:02N Null 26 30

DRB5*01:01:01-01:14, DRB5*02:02-02:06

DR51, 

Null, -

31

Length of spec. PCR product(s)

4

210200200210120

8085

100210170

90

175100

85

215195175100110110160215175

260 220 220210210175235215135 170105 225 140135175240

250180 155170225

Well No.

ser

5

9 1011121314151617181920212223242526272829303132

Specific Reaction

1st match: DRB*

2nd match: DRB*

Control Reaction

Failed Controls

False Positive

False Negative

Negative control



Neg. contr.


1The DRB1*03:11:02 allele has been renamed DRB1*03:81.
The sequence of the DRB1*04:05:12 allele has been shown to be identical to DRB1*04:05:14.
The sequence of the DRB1*0702 allele has been shown to be identical to DRB1*07:01:01:01.
The sequence of the DRB1*08031 allele has been shown to be identical to DRB1*08:03:02.

The sequence of the DRB1*09011 allele has been shown to be identical to DRB1*09:01:02.

The sequence of the DRB1*1171 allele has been shown to be identical to DRB1*11:02:01.

The sequence of the DRB1*12031 allele has been shown to be identical to DRB1*12:01:01.

The DRB1*1466 allele has been renamed DRB1*14:32:02.

The sequence of the DRB1*1606 allele has been shown to be identical to DRB1*16:05:01.

The sequence of the DRB3*010101 allele has been shown to be identical to DRB3*01:01:02:01.
The DRB4*0101102N allele has been renamed DRB4*01:03:10:02N.

The sequence of the DRB5*0201 allele has been shown to be identical to DRB5*02:02.

Due to sharing of sequence motifs in codon 38, DRB3*01:14 will also be amplified in primer mixes 13, 14 and 25 in addition to primer mix 29.

2The DRB1*08:09 and the DRB1*14:15 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.
The DRB1*08:20 and the DRB1*13:18, 13:47 and 13:55 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*08:31, 08:41 and DRB1*11:67 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

The DRB1*13:13 and 13:119 and the DRB1*14:84 and 14:116 alleles yield identical amplification patterns with the DR low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

3The serological reactivity of all DRB alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele. 

4Primer mix 32 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 430 base pairs.

 ‘ser’, serological HLA specificity. 
‘w’, may be weakly amplified.
Change in revision R01 compared to R00:

1. The serological HLA specificities have been corrected in the Interpretation tables.







Olerup SSP AB

For In Vitro Diagnostic Use 
Franzengatan 5

www.olerup-ssp.com
SE-112 51 Stockholm
Sweden

