DQA1*02,05;DQB1*02,03:02	Lot No: 3N9	Expiry Date: 2025-06-01

(101.903-24/24u)

Name:___________________

Sample ID:_______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:________________

Reviewed By:________________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

The primer pair in well 19 will in some samples give rise to two HLA-specific PCR fragments and may give rise to a lower yield for the DQB1*03xx alleles.

Primer mixes 12 and 13 may have tendencies of unspecific amplification, most pronounced in primer mix 12.

Primer mix 8 may give rise to a lower yield of HLA-specific PCR product than the other DQA1*02,05;DQB1*02,03:02 primer mixes.

Primer mix 23 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1DQA1 and DQB1 alleles listed on the IMGT/HLA web page 2021-April-20, release 3.44.0, www.ebi.ac.uk/imgt/hla.

Bold lettering alleles are DQA1*02,05;DQB1*02,03:02 associated alleles.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

Abbreviations

?: nucleotide sequence information not available for the primer matching sequence.

Expected Results

The table below describes expected results for the (groups of) alleles that the kit is able to detect and separate.

Table 1: Expected results for targeted DQA1 and DQB1 alleles.

| DQA1 alleles | DQB1 alleles | Positive DQA1 wells | Positive DQB1 wells |
| --- | --- | --- | --- |
| 05:01 | 02:01 (DQ2) | 3, 5, 11 | 12, 14, 16 |
| (02:01) 05:05 | 02:02 (DQ2) (03:01) | 1, 8 3, 5, 6, 10 | 12, 13, 16 14, 19, 21 |
| 03:01 | 03:02 (DQ8) | 2, 8 | 14, 16, 17, 19, 22 |
| 03:02, 03:03 | 03:02 (DQ8) | 2 | 14, 16, 17, 19, 22 |
| 05:01 |  | 3, 5, 11 |  |
| 05:05 |  | 3, 5, 6, 10 |  |
| 02:01 |  | 1, 8 |  |
| 03:01 |  | 2, 8 |  |
| 03:02, 03:03 |  | 2 |  |
|  | 02:01 |  | 12, 14, 16 |
|  | 02:02 |  | 12, 13, 16 |
|  | 03:01 |  | 14, 19, 21 |
|  | 03:02 |  | 14, 16, 17, 19, 22 |

The negative control DNA must only give rise to the internal control bands of 430 or 515 base pairs respectively and no DQA1*02,05;DQB1*02,03:02 specific bands. Additional bands might indicate inappropriate test conditions or contamination.