DQB1*02 (101.213-24/24u) 	Lot No: 8N0	Expiry Date: 2026-11-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 2 and 3 may have tendencies of unspecific amplifications.

Primer mix 32 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1HLA-DQB1 in bold lettering are listed as confirmed alleles on the IMGT/HLA web page 2016-October-14, release 3.26.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following DQB1*02 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified DQB1*02 alleles | Other amplified DQB1 alleles |
| --- | --- | --- | --- |
| 5 | 95 bp 140 bp 195 bp | *02:23 *02:04, 02:37 *02:96N |  |
| 9 | 95 bp 195 bp | *02:16 *02:07:01-02:07:02 |  |
| 10 | 180 bp 230 bp | *02:08 *02:12 |  |
| 11 | 95 bp 170 bp | *02:24 *02:09 | *03:330 |
| 12 | 120 bp 160 bp | *02:11, 02:25 *02:13, 02:30 |  |
| 13 | 120 bp 215 bp | *02:17, 02:32 *02:38 |  |
| 14 | 145 bp 220 bp | *02:18N, 02:54 *02:34, 02:38 |  |
| 15 | 100 bp 205 bp | *02:21, 02:39 *02:35 |  |
| 16 | 100 bp 150 bp 230 bp | *02:22, 02:39 *02:54 *02:20N |  |
| 19 | 100 bp 165 bp | *02:50 *02:33, 02:36 | *03:479 |
| 21 | 100 bp 130 bp | *02:23, 02:31 *02:40 |  |
| 22 | 90 bp 175 bp | *02:27 *02:28 |  |
| 23 | 100 bp 255 bp | *02:53Q *02:41 | *03:322 |
| 24 | 160 bp 200 bp | *02:51 *02:42 |  |
| 25 | 160 bp 195 bp 220 bp | *02:51 *02:63 *02:46 |  |
| 27 | 155 bp 210 bp | *02:62 *02:72 |  |
| 28 | 150 bp 195 bp | *02:64, 02:79 *02:96N | *03:21, 03:452, 05:35, 05:182, 05:259, 06:243, 06:255 |
| 29 | 65 bp 255 bp | *02:67N *02:63 | *03:48, 03:456 |
| 30 | 110 bp 185 bp | *02:58N *02:29 |  |

4The following DQB1*02 alleles can be distinguished by the different sizes of the specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| DQB1*02:18N, 02:34 | 14 |
| DQB1*02:21, 02:35 | 15 |
| DQB1*02:29, 02:58N | 30 |