# Gel Documentation Form and Worksheet

# DQB1*05 (101.211-24/24u) 	Lot No: 8G3	Expiry Date: 2021-03-01

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

## PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 15 may have tendency of unspecific amplification.

Primer mixes 26 and 30 may give rise to a lower yield of HLA-specific PCR product than the other DQB1*05 primer mixes.

Primer mix 32 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1DQB1*05 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.26.0, October 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following DQB1*05 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified DQB1*05 alleles | Other amplified DQB1 alleles |
| --- | --- | --- | --- |
| 5 | 120 bp 185 bp | *05:04, 05:52, 05:77, 05:132Q, 05:146 *05:10 |  |
| 7 | 185 bp 245 bp 270 bp | *05:06:01-05:06:02, 05:07, 05:50 *05:52 *05:112 |  |
| 8 | 135 bp 190 bp | *05:27, 05:87Q, 05:132Q *05:09, 05:29 |  |
| 10 | 115 bp 195 bp | *05:23 *05:12, 05:71 |  |
| 11 | 100 bp 150 bp 190 bp | *05:13, 05:32, 05:42 , 05:45 , 05:117 *05:35 *05:96 | *03:196 *02:64, 02:79, 03:21, 06:243 |
| 12 | 120 bp 150 bp 195 bp | *05:25 *05:14 *05:41N, 05:90N, 05:96 | *02:96N |
| 13 | 145 bp 180 bp 220 bp | *05:40 *05:15, 05:33, 05:49 *05:34, 05:44, 05:128N, 05:130 |  |
| 14 | 155 bp 205 bp | *05:40, 05:69, 05:101, 05:145, 05:153 *05:16, 05:73, 05:98, 05:116 | *06:156, 06:162, 06:169 |
| 15 | 65 bp 105 bp 135 bp | *05:03:06, 05:03:14 *05:17 *05:37, 05:104 | *03:03:05 *06:250 |
| 16 | 145 bp 195 bp | *05:20, 05:105 *05:47, 05:159 | *02:99, 03:249, 06:105, 06:185, 06:225 |
| 17 | 125 bp 200 bp | *05:25, 05:31, 05:46, 05:108 *05:41N, 05:90N, 05:106 | *06:111 *02:96N |
| 18 | 180 bp 220 bp | *05:24 *05:72, 05:128N |  |
| 19 | 195 bp 225 bp | *05:28, 05:30 *05:21, 05:60 |  |
| 20 | 110 bp 210 bp 270 bp | *05:32, 05:42, 05:57 *05:106 *05:38, 05:62, 05:119 | *06:146:01 |
| 21 | 140 bp 165 bp 200 bp 230 bp | *05:37, 05:88, 05:104 *05:63 *05:39 *05:26, 05:82 | *06:250 |
| 24 | 150 bp 185 bp | *05:71 *05:73, 05:80, 05:98, 05:116 | *06:28, 06:56, 06:79:01-06:79:02, 06:89 |
| 25 | 115 bp 180 bp | *05:75, 05:111 *05:112 |  |
| 26 | 85 bp 185 bp | *05:103 *05:78 |  |
| 27 | 145 bp 200 bp 280 bp | *05:88 *05:110N *05:79, 05:156 | *06:158N |
| 28 | 135 bp 170 bp | *05:56 *05:107 |  |
| 30 | 100 bp 145 bp | *05:105, 05:111 *05:51 | *03:03:05, 06:02:19, 06:03:08, 06:79:02 |

4The following DQB1*05 alleles can be distinguished by the different sizes of the specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| DQB1*05:16, 05:101 | 14 | DQB1*05:35, 05:117 | 11 |
| DQB1*05:20, 05:159 | 16 | DQB1*05:110N, 05:156 | 27 |
| DQB1*05:27, 05:29 | 8 |

’w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

# Change in revision R01 compared to R00:

# Primer mix 12 does not amplify the DQB1*05:84 and DQB1*06:103 alleles. This correction has been implemented in the Specificity and Interpretation tables.