DQB1*05 (101.211-24/24u) 	Lot No: 4S6	Expiry Date: 2027-09-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 15 and 16 may have a tendency of unspecific amplification.

Primer mixes 26 and 30 may give rise to a lower yield of HLA-specific PCR product than the other DQB1*05 primer mixes.

Primer mix 17 may have a tendency to give rise to primer oligomer formation.

Primer mix 32 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1DQB1*05 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.26.0, October 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following DQB1*05 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified DQB1*05 alleles | Other amplified DQB1 alleles |
| --- | --- | --- | --- |
| 2 | 135 bp 225 bp | *05:01:01:01-05:01:48, 05:07, 05:11:01-05:12, 05:18:01-05:18:02, 05:20:01-05:20:02, 05:22, 05:25, 05:27, 05:29-05:32, 05:44-05:45, 05:48-05:49, 05:51, 05:54-05:55, 05:61-05:63, 05:68-05:76, 05:80-05:81, 05:84, 05:88-05:89:02, 05:92-05:93, 05:95, 05:99, 05:103-05:104, 05:107, 05:110N-05:112, 05:114-05:115, 05:120, 05:122, 05:124, 05:126-05:128N, 05:133, 05:137-05:139, 05:141, 05:144, 05:148, 05:150-05:152, 05:154-05:160, 05:162-05:164, 05:166-05:169, 05:171, 05:173, 05:176-05:177, 05:180, 05:182-05:185N, 05:187-05:188, 05:190, 05:193-05:195, 05:197, 05:215N-05:217, 05:219, 05:223, 05:225-05:226, 05:228, 05:230, 05:232, 05:234, 05:237, 05:240, 05:242, 05:246, 05:248-05:249, 05:252, 05:255-05:256, 05:258, 05:261, 05:263, 05:266-05:272, 05:275, 05:277, 05:279, 05:283N, 05:285-05:286, 05:288, 05:290-05:292, 05:295N, 05:297-05:300, 05:302-05:303, 05:307-05:308, 05:310, 05:313Q, 05:315-05:317, 05:319, 05:322-05:323 *05:218 | *06:325, 06:389 |
| 5 | 120 bp 185 bp | *05:04, 05:52, 05:77, 05:132Q, 05:146, 05:262 *05:10 | *03:309:01 |
| 7 | 180 bp 245 bp 270 bp | *05:06:01-05:06:02, 05:07, 05:50 *05:20:02, 05:52 *05:112 |  |
| 8 | 135 bp 190 bp | *05:27, 05:87Q, 05:132Q *05:09, 05:29 | *03:420 |
| 10 | 115 bp 195 bp | *05:23, 05:242 *05:12, 05:71 |  |
| 11 | 100 bp 150 bp 190 bp | *05:13, 05:32, 05:42, 05:45, 05:117, 05:169, 05:201 *05:35, 05:182, 05:259 *05:96 | *04:56 *02:64, 02:79, 03:21, 03:452, 06:243, 06:255 |
| 12 | 120 bp 150 bp 195 bp | *05:25 *05:14, 05:84, 05:211 *05:41N, 05:90N, 05:96 | *02:169, 03:319, 03:347, 04:70, 06:103, 06:284, 06:350 *02:96N |
| 13 | 145 bp 180 bp 220 bp | *05:40 *05:15, 05:33, 05:49, 05:321 *05:34, 05:44, 05:128N, 05:130 |  |
| 14 | 120 bp 155 bp 205 bp | *05:58 *05:40, 05:69, 05:101, 05:145, 05:153 *05:16, 05:73, 05:98, 05:116 | *06:156, 06:162, 06:169 |
| 15 | 65 bp 105 bp 135 bp | *05:03:06, 05:03:14, 05:03:20 *05:17 *05:37, 05:104 | *03:03:05 *06:250 |
| 16 | 145 bp 195 bp | *05:20:01, 05:105 *05:47, 05:159, 05:312 | *02:99, 03:249, 03:312, 03:481, 06:105, 06:185, 06:225, 06:407, 06:432 |
| 17 | 125 bp 200 bp | *05:25, 05:31, 05:46, 05:108 *05:41N, 05:90N, 05:106 | *03:317:01-03:317:02, 06:111, 06:265, 06:455 *02:96N |
| 18 | 120 bp 180 bp 220 bp | *05:58 *05:24 *05:72, 05:80, 05:128N | *06:325 |
| 19 | 195 bp 225 bp | *05:28, 05:30 *05:21, 05:60 |  |
| 20 | 110 bp 210 bp 270 bp | *05:32, 05:42, 05:57 *05:106 *05:38, 05:62, 05:119 | *03:432, 04:56 *06:146:01 |
| 21 | 140 bp 165 bp 200 bp 230 bp | *05:37, 05:88, 05:104 *05:63 *05:39 *05:26, 05:82, 05:302 | *06:250 |
| 24 | 150 bp 190 bp | *05:71 *05:73, 05:80, 05:98, 05:116 | *06:28, 06:56, 06:79:01-06:79:02, 06:89, 06:451 |
| 25 | 115 bp 180 bp | *05:75, 05:111 *05:112 |  |
| 26 | 85 bp 185 bp | *05:103, 05:231 *05:78 |  |
| 27 | 145 bp 200 bp 280 bp | *05:88 *05:110N, 05:236N *05:79, 05:156 | *06:158N |
| 28 | 135 bp 170 bp | *05:56 *05:107 |  |
| 30 | 100 bp 145 bp | *05:105, 05:111 *05:51 | *03:03:05, 06:02:19, 06:03:08, 06:09:11, 06:13:03, 06:79:02 |
| 31 | 125 bp 150 bp 220 bp | *05:102, 05:308 *05:67 *05:54 | *02:157, 03:190, 03:506, 04:42, 06:223, 06:406 |

4The following DQB1*05 alleles can be distinguished by the different sizes of the specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| DQB1*05:12, 05:242 | 10 | DQB1*05:27, 05:29 | 8 |
| DQB1*05:13, 05:201, 05:259 | 11 | DQB1*05:35, 05:117 | 11 |
| DQB1*05:16, 05:101 | 14 | DQB1*05:45, 05:169, 05:182 | 11 |
| DQB1*05:20:01, 05:159 | 16 | DQB1*05:110N, 05:156 | 27 |

5The following alleles will give rise to identical amplification patterns with the DQB1*05 subtyping kit. These alleles can be distinguished by the DQ low resolution kit and/or the DQB1*06 subtyping kit:

Alleles

DQB1*05:300, DQB1*06:389

6This lot of the DQB1*05 kit cannot separate the DQB1*05:19 and 05:97 alleles.

Abbreviations

w: might be weakly amplified.

?: nucleotide sequence information not available for the primer matching sequence.

Changes in revision R01 compared to R00:

- Primer mixes 3 and 29 amplifies the DQB1*05:97 allele. This correction has been implemented in the Specificity and Interpretation tables.