# Gel Documentation Form and Worksheet

# DRB1*04 (101.114-24/03, -24u/03u) 	Lot No: 4K2	Expiry Date: 2023-10-01

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

# Interpretation:___________                            Failed lanes: ____________                            Comments:________________

## PHOTO DOCUMENT

Gel Picture

Abbreviations

ICB: Internal Control Band,

AmpS: Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 21, 22, 29 and 31 may have tendencies of unspecific amplifications.

Primer mixes 14 and 31 have a tendency to giving rise to primer oligomer formation.

Primer mix 48 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1DRB1*04 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.27.0, January 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following DRB1*04 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | #### Amplified DRB1*04 alleles | Other amplified DRB alleles |
| --- | --- | --- | --- |
| 11 | 175 bp 220 bp | *04:15, 04:61, 04:138, 04:145, 04:154, 04:170, 04:265, 04:291 *04:22, 04:98:02, 04:283 | *11:22 |
| 12 | 180 bp 265 bp | *04:16, 04:180 *04:23, 04:52, 04:151, 04:203, 04:207 |  |
| 13 | 170 bp 225 bp | *04:46 *04:12, 04:18, 04:25, 04:31, 04:54-04:55, 04:58, 04:86, 04:103, 04:126, 04:185, 04:191 |  |
| 15 | 175 bp 215 bp | *04:123 *04:34 | *01:10 |
| 19 | 140 bp 210 bp | *04:39 *04:24, 04:69, 04:78, 04:101, 04:168, 04:251 | *14:10, 14:143, 14:182 |
| 20 | 100 bp 175 bp 255 bp | *04:63, 04:110, 04:161, 04:253, 04:264N *04:40 *04:27 |  |
| 21 | 155 bp 210 bp | *04:29 *04:37, 04:58, 04:73:01-04:73:02, 04:105:01-04:105:02, 04:122, 04:132, 04:206, 04:284 |  |
| 23 | 135 bp 255 bp | *04:33 *04:201 |  |
| 25 | 125 bp 230 bp 260 bp | *04:26 *04:45 *04:64 |  |
| 27 | 115 bp 265 bp | *04:10:01-04:11:04, 04:11:06-04:12, 04:60, 04:67, 04:91, 04:126, 04:136, 04:147, 04:201, 04:203, 04:207, 04:219, 04:250 *04:49 | *03:12, 03:158, 08:06, 08:10, 08:12, 08:22, 08:54, 08:70, 12:46, 13:04, 13:32, 13:48, 13:58, 13:75, 13:81, 13:89:01-13:89:02, 13:93-13:94:02, 13:108, 13:164, 13:169, 13:276, 14:65, 14:78, 14:169-14:170, 15:12 |
| 29 | 75 bp 270 bp | *04:07:01:01-04:07:06, 04:17:01-04:17:02, 04:20, 04:69, 04:74, 04:92, 04:102, 04:129, 04:160, 04:170, 04:172, 04:197, 04:199, 04:209, 04:212N, 04:222, 04:246, 04:249, 04:256, 04:262-04:263, 04:269, 04:278 *04:19?-04:21?, 04:64, 04:192, 04:213 | *01:17, 09:01:02:01-09:01:05, 09:01:08-09:02:02, 09:04-09:05, 09:07-09:08, 09:10-09:41, 10:29, 14:07:01-14:07:02, 14:14, 14:36, 14:42, 14:44:01-14:44:03, 14:51, 14:68:01-14:68:02, 14:93, 14:134, 14:143, DRB4*01:60, DRB5*01:12 DRB3*02:84 |
| 31 | 85 bp 200 bp 260 bp | *04:14, 04:106, 04:132 *04:214N *04:64 | *01:03:01-01:03:04, 01:39N, 01:89, 01:99, 01:102, 08:40, 08:92, 11:11:01, 11:11:03, 11:14:01-11:14:02, 11:20, 11:68, 11:73, 11:93, 11:122, 11:168, 11:184, 11:237, 13:02:01:01-13:02:03, 13:02:05-13:02:09, 13:02:11-13:03:09, 13:23:01-13:23:02, 13:29, 13:31, 13:33:01-13:34, 13:36-13:39, 13:41, 13:45, 13:63, 13:65-13:67, 13:73-13:74, 13:85, 13:88, 13:90, 13:95-13:97:02, 13:99, 13:101, 13:103-13:104, 13:107, 13:115, 13:120, 13:122-13:124, 13:126, 13:128, 13:135, 13:139, 13:143, 13:145, 13:147, 13:151-13:152, 13:155, 13:159, 13:165, 13:167-13:168, 13:170-13:172, 13:174, 13:179-13:180, 13:182, 13:188, 13:194-13:195, 13:198, 13:202, 13:207-13:212, 13:216, 13:219-13:220, 13:225, 13:227, 13:230, 13:236-13:237, 13:239-13:241, 13:243-13:244, 13:247, 13:250, 13:253, 13:255N, 13:257, 13:264-13:265, 13:268N-13:269, 13:275, 13:280, 13:282, 13:285 |
| 35 | 210 bp 245 bp | *04:81N *04:51, 04:70, 04:260 |  |
| 36 | 120 bp 180 bp | *04:96 *04:71 | *03:149 |
| 38 | 140 bp 225 bp | *04:89 *04:73:01-04:73:02, 04:75, 04:105:01-04:105:02, 04:122, 04:178N, 04:284 |  |
| 39 | 100 bp 175 bp | *04:83 *04:15, 04:61-04:62, 04:69, 04:138, 04:145-04:146, 04:154, 04:170, 04:235, 04:251, 04:265, 04:291 | *11:22, 14:10, 14:57, 14:143, 14:182 |
| 40 | 145 bp 200 bp | *04:30, 04:85 *04:38, 04:56:01-04:56:02, 04:88, 04:100, 04:107, 04:125, 04:205 |  |
| 41 | 165 bp 205 bp | *04:76, 04:108 *04:182 |  |
| 42 | 110 bp 140 bp | *04:41, 04:77, 04:102, 04:113, 04:152, 04:164, 04:211, 04:225, 04:230, 04:250 *04:79 | *14:10, 14:57, 14:143, 14:182 |
| 44 | 165 bp 240 bp | *04:127 *04:87, 04:94:01N, 04:120N, 04:129, 04:157N, 04:186N, 04:212N |  |
| 46 | 75 bp 200 bp | *04:92, 04:274 *04:214N |  |
| 47 | 140 bp 250 bp | *04:117 *04:119N, 04:142N |  |

4The following DRB1*04 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| DRB1*04:16, 04:151 | 12 | DRB1*04:49, 04:219 | 27 |
| DRB1*04:23, 04:180 | 12 | DRB1*04:79, 04:211 | 42 |
| DRB1*04:27, 04:110 | 20 | DRB1*04:89, 04:178N | 38 |
| DRB1*04:39, 04:101 | 19 | DRB1*04:117, 04:119N | 47 |

Abbreviations

w: might be weakly amplified.

?: nucleotide sequence information not available for the primer matching sequence.