Gel Documentation Form and Worksheet

DRB1*08 (101.127-12/04, -12u/04u) 	Lot No: 0L6	Expiry Date: 2024-04-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________	Failed lanes: ___________ 	 Comments:_________________

Gel Picture

PHOTO DOCUMENT

Abbreviations

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 2 and 5 may have tendencies of unspecific amplifications.

Primer mixes 3 and 24 might have a tendency to giving rise to primer oligomer formation.

Primer mix 24 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by control primer pairs. PCR product sizes range from 75 to 200 base pairs. The PCR product generated by the control primer pair is 200 base pairs.

1DRB1*08 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.31.0, January 2018.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following DRB1*08 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified DRB1*08 alleles | Other amplified DRB1 alleles |
| --- | --- | --- | --- |
| 9 | 125 bp 175 bp | *08:26, 08:35-08:36:02 *08:14 | *03:12, 03:158, 04:202, 13:32, 13:65, 13:93, 13:120, 13:139, 14:13, 14:63, 14:65, 14:78, 14:85, 14:169-14:170 |
| 10 | 205 bp 250 bp | *08:40, 08:92 *08:12, 08:22 | *11:229, 12:73, 13:17, 13:116, 13:175, 13:214 *12:01:01:01-12:02:10, 12:04-12:15, 12:17-12:18, 12:20-12:21, 12:23-12:38, 12:40-12:41, 12:43-12:56, 12:58-12:82, 14:28 |
| 12 | 95 bp 145 bp 195 bp 220 bp | *08:27 *08:44 *08:13, 08:48 *08:43 | *12:12, 12:62 |
| 13 | 135 bp 165 bp 260 bp | *08:33 *08:11 *08:39 |  |
| 14 | 135 bp 250 bp | *08:17, 08:28, 08:37, 08:45:01-08:45:02, 08:63 *08:60N | *03:92, 11:67, 11:193:01-11:193:02, 11:209, 11:229, 12:01:01:01-12:20, 12:22-12:37, 12:39-12:41, 12:43-12:52, 12:54-12:56, 12:58-12:75, 12:77-12:82, 13:17, 13:175, 13:214, 14:138, 14:175 |
| 17 | 75 bp 115 bp 265 bp | *08:04:01, 08:04:04-08:04:07, 08:06, 08:10, 08:12, 08:20, 08:22, 08:28, 08:54, 08:59, 08:67, 08:70, 08:75, 08:80, 08:83, 08:94, 08:98-08:99 *08:14 | *04:12, 04:18, 04:25, 04:58, 11:25, 11:67, 11:119, 13:18, 13:144, 13:156, 13:164, 14:12:01-14:12:02, 14:15, 14:78, 14:84, 14:156, 14:173, 15:21 *11:261 *11:261 |
| 18 | 150 bp 225 bp | *08:09:01-08:09:02, 08:21, *08:32, 08:35, 08:87 *08:20 | *13:232, 14:15, 14:40, 14:55, 14:77, 14:84, 14:136, 14:164, 14:173 *03:157, 11:23:01-11:23:02, 11:25, 11:45, 11:55, 11:64, 11:119, 13:13, 13:18, 13:47, 13:55, 13:119, 13:144, 13:146, 13:154, 13:156, 13:158, 13:164, 13:232, 13:286, 14:03:01-14:03:02, 14:12:01-14:12:02, 14:27:01-14:27:02, 14:40, 14:55, 14:63, 14:67, 14:77-14:78, 14:84-14:85, 14:89, 14:102, 14:115-14:116, 14:135-14:136, 14:144, 14:156, 14:164, 14:173-14:174, 14:181, 14:218 |
| 19 | 100 bp 170 bp | *08:16, 08:38 *08:07, 08:49 |  |
| 21 | 70 bp 120 bp | *08:78N *08:17, 08:28, 08:37, 08:45:01-08:45:02, 08:63 | *03:157, 11:23:01-11:23:02, 11:25, 11:45, 11:55, 11:64, 11:67, 11:119, 13:18, 13:119, 13:144, 13:146, 13:154, 13:156, 13:158, 13:286, 15:21 |
| 22 | 130 bp 165 bp 215 bp | *08:23 *08:19 *08:29 | *12:66, 14:04:01:01-14:04:07, 14:11, 14:28, 14:68:01-14:68:02, 14:71, 14:73, 14:93, 14:120, 14:126:01-14:126:02, 14:138, 14:145, 14:148, 14:152N, 14:175, 14:180, 14:196, 14:201, 14:205-14:206, 14:210Q |
| 23 | 220 bp 250 bp | *08:43 *08:04:01, 08:04:04-08:04:07, 08:06, 08:10, 08:28, 08:31, 08:54, 08:59, 08:67, 08:70, 08:75, 08:83, 08:94, 08:98-08:99 | *03:92, 11:67, 11:209, 11:229, 12:03:02-12:03:03, 12:19, 13:17, 13:116, 13:175, 13:214, 14:04:01:01-14:04:07, 14:11, 14:15, 14:31, 14:50, 14:52, 14:73, 14:76, 14:79, 14:107, 14:120, 14:126:01-14:126:02, 14:138, 14:145, 14:148, 14:152N, 14:175, 14:180, 14:196, 14:201, 14:205-14:206, 14:210Q |

4The following alleles cannot be separated by the HLA-DRB1*08 primer set. These alleles can be distinguished by the DR low resolution kit and/or the DRB1*14 high resolution kit.

Alleles

DRB1*08:20, DRB1*14:12:01-14:12:02, 14:84, 14:156, 14:173

Abbreviations

‘w’, may be weakly amplified.