PAGE  
January 2010
DRB1*10 
Page 1 of 1
Rev. No: 00
79G
Worksheet can be downloaded from

Worksheet
www.olerup-ssp.com








	PHOTO DOCUMENT


Olerup SSP®
Olerup SSP® DRB1*10
Prod. No:
101.129-06/06u
Lot No:    
79G
Expiry Date:    
2012-January-01
Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 4 primer mixes used for DRB1*10 SSP typing.
	Primer Mix
	Size of spec. PCR product1
	Size of control band2
	Amplified DRB1*10 alleles

	  1
	210 bp
	515 bp
	100101-1003

	  2
	205 bp
	515 bp
	100101-100103, 1003

	  3
	205 bp
	430 bp
	1002

	  43
	  70 bp
	430 bp
	1003


1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*10 SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given. 

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Wells number 1 and 2 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*10 subtyping and in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Specific PCR fragments shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR bands.

[image: image1.emf]Well No. 1 2 3 4

DRB1 allele

*100101-100103 1 2

*1002 1 3

*1003 1 2 4

Length of spec. PCR product

210 205 205

70

Well No. 1 2 3 4

Specific Reaction

1st match: DRB1*

2nd match: DRB1*

Control Reaction

Failed Controls

False Positive

False Negative








Olerup SSP AB

For In Vitro Diagnostic Use
Hasselstigen 1

www.olerup-ssp.com
SE-133 33 Saltsjöbaden

Sweden

