PAGE  
March 2008
DRB1*12 
Page 1 of 3
Rev. No: 00
25E
Worksheet can be downloaded from

Worksheet
www.olerup.com









	PHOTO DOCUMENT


Olerup SSP®
Olerup SSP® DRB1*12
Prod. No:
101.128-12 – including Taq polymerase

101.128-12u – without Taq polymerase
Lot No:    
25E
Expiry Date:    
2010-March-01
Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 22 primer mixes used for DRB1*12 SSP typing.
	Primer Mix
	Size of spec. PCR product1
	Size of control band2
	Amplified DRB1*12 alleles3
	Other amplified DRB1 alleles4

	  1
	135 bp
	515 bp
	120101-120202, 120302-1216
	0817, 0828, 1167, 1317

	  2
	215 bp
	430 bp
	120101-120202, 120302-1216
	0805, 0818, 0824, 0825, 0831, 1317, 1431, 1452

	  3
	165 bp
	430 bp
	1209
	080201-080203, 080401-080404, 0809, 0813, 0821, 0824, 0828, 0830, 1317, 1415, 1452

	  4
	105 bp
	430 bp
	120101-120202, 120302, 1204, 1206-1213, 1216
	0832

	  5
	165 bp
	515 bp
	120101-120202, 120302, 1205-1208, 1210-1216
	0819, 0825, 0834

	  6
	250 bp
	430 bp
	120101-120202, 1204-1215
	0812, 0822, 1428

	  7
	215 bp
	430 bp
	120101-120202, 1204-1207, 1209-1212, 1213weakly, 1214, 1215
	

	  8
	195 bp
	430 bp
	120101-120102, 120302-1206, 1208-1211, 1214
	080302, 0810, 0812, 0814, 0815, 0818, 0819, 0823, 0825, 0827, 0829, 0830, 0832-0834,  1317

	
	
	
	
	

	  9
	165  bp
	430 bp
	120101, 120201, 1204, 1206-1213
	

	10
	195 bp
	430 bp
	120201-120202, 1213, 1215, 1216
	080101-080203, 080401-0809, 0811, 0816, 0817, 0821, 0822, 0824, 0826, 0828, 0831, 1167, 1415, 1473

	11
	250 bp
	430 bp
	120302
	080401, 080402-080403weakly, 080404, 0806, 0810, 0828, 0831, 1167, 1317, 1404, 1411, 1415, 1431, 1450, 1452, 1473

	12
	 170 bp
	515 bp
	1204
	0831, 1167, 1411

	13
	105 bp
	515 bp
	1205, 1214, 1215
	0809, 0821, 1404, 1411, 1415, 1428, 1431, 1450

	14
	135 bp
	515 bp
	1206
	150101-1527weakly, 160101-160502weakly, 1607-1613Nweakly

	15
	200 bp
	515 bp
	1207
	

	16
	  75 bp
	430 bp
	1208
	1334, 1364, 1441

	
	
	
	
	

	17
	  80 bp
	430 bp
	1210
	

	18
	135 bp
	515 bp
	1211
	

	19
	195 bp
	430 bp
	1212
	0813

	20
	220 bp
	430 bp
	1213
	

	21
	105 bp
	430 bp
	120101-120202, 120302-1213, 1215, 1216
	0832

	22
	220
	430 bp
	1216
	0832


1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*12 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes are given. 

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*12 subtyping.

In addition, wells number 5, 12 to 15 and 18 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB alleles only partial second exon and fourth exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*0101 consensus sequence. 

4Due to the sharing of sequence motifs within the DR52 group of DRB1 alleles, most DRB1*08, one DRB1*11, three DRB1*13 and nine DRB1*14 alleles will be amplified by some of the DRB1*12 primer mixes. In addition, the DRB1*15 and DRB1*16 alleles will be weakly amplified by primer mix 14.

[image: image1.emf]Well No.

1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22

DRB1 allele

1 

*120101

1 2 4 5 6 7 8 9 21

*120102

1 2 4 5 6 7 8 21

*120201

1 2 4 5 6 7 9 10 21

*120202

1 2 4 5 6 7 10 21

*120302

1 2 4 5 8 11 21

*1204

1 2 4 6 7 8 9 12 21

*1205

1 2 5 6 7 8 13 21

*1206

1 2 4 5 6 7 8 9 14 21

*1207

1 2 4 5 6 7 9 15 21

*1208

1 2 4 5 6 8 9 16 21

*1209

1 2 3 4 6 7 8 9 21

*1210

1 2 4 5 6 7 8 9 17 21

*1211

1 2 4 5 6 7 8 9 18 21

*1212

1 2 4 5 6 7 9 19 21

*1213

1 2 4 5 6 w 9 10 20 21

*1214

1 2 5 6 7 8 13

*1215

1 2 5 6 7 10 13 21

*1216

1 2 4 5 10 21 22

*080101-080103, 0807, 0808, 0811, 

0816, 0826

10

*080201-080203

3 10

*080302, 0814, 0815, 0823, 0827, 

0829, 0833

8

*080401, 080404

3 10 11



[image: image2.emf]*080402-080403

3 10 w

*0805

2 10

*0806, 1473

10 11

*0809, 0821

3 10 13

*0810

8 11

*0812

6 8

*0813

3 19

*0817

1 10

*0818

2 8

*0819, 0834

5 8

*0822

6 10

*0824

2 3 10

*0825

2 5 8

*0828

1 3 10 11

*0830

3 8

*0831

2 10 11 12

*0832

4 8 21 22

*1167

1 10 11 12

*1317

1 2 3 8 11

*1334, 1364, 1441

16

*1404, 1450

11 13

*1411

11 12 13

*1415

3 10 11 13

*1428

6 13

*1431

2 11 13

*1452

2 3 11

*150101-1527, 160101-160502, 

1607-1613N

w

Length of spec. PCR product

135 215 165 105 165 250 215 195 165 195 250 170 105 135 200

75 80

135 195 220 105 220

Well No.

1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22

Specific Reaction

1st match: DRB1*

2nd match: DRB1*

Control Reaction

Failed Controls

False Positive

False Negative

4The sequence of the DRB1*120301 allele has been shown to be identical to DRB1*1201.
‘w’, may be weakly amplified.







Olerup SSP AB
In U.S.: For Research Use Only.
www.olerup.com
Hasselstigen 1
Not For Use in Diagnostic Procedures.
SE-133 33 Saltsjöbaden

Sweden

