PAGE  
October 2011
DRB1*12 
Page 2 of 4
Rev. No: 00
83M
Worksheet can be downloaded from


Worksheet
www.olerup-ssp.com





	PHOTO DOCUMENT


Olerup SSP®
Olerup SSP® DRB1*12
Prod. No:
101.128-12/12u

Lot No:    
83M
Expiry Date:    
2014-April-01
Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for DRB1*12 SSP typing.
	Primer Mix
	Size of spec. PCR product1
	Size of control band2
	Amplified DRB1*12 alleles3
	Other amplified DRB1 alleles4

	  1
	135 bp
	515 bp
	*12:01:01-12:20, 12:22-12:27
	*08:17, 08:28, 08:37, 08:45, 11:67, 13:17

	  2
	215 bp
	430 bp
	*12:01:01-12:21, 12:23-12:27
	*08:05, 08:18, 08:24-08:25, 08:31, 08:40-08:41, 13:17, 13:116, 14:31, 14:52

	  3
	165 bp
	430 bp
	*12:09
	*08:02:01-08:02:03, 08:04:01-08:04:05, 08:04:07, 08:09, 08:13, 08:21, 08:24, 08:28, 08:30:01, 08:42, 08:44-08:45, 13:17, 13:116, 14:15, 14:52

	  45
	105 bp
	430 bp
	*12:01:01-12:04, 12:06-12:13, 12:16-12:27
	*08:32

	  5
	165 bp
	515 bp
	*12:01:01-12:03:02, 12:05-12:08, 12:10-12:17, 12:19-12:23, 12:24Nw, 12:25-12:27
	*08:19, 08:25, 08:34

	  6
	250 bp
	430 bp
	*12:01:01-12:02:05, 12:04-12:15, 12:17-12:18, 12:20-12:21, 12:23-12:27
	*08:12, 08:22, 14:28

	  7
	215 bp
	430 bp
	*12:01:01-12:02:05, 12:04-12:07, 12:09-12:12, 12:13w, 12:14-12:15, 12:17-12:18, 12:20-12:21, 12:24N-12:26
	

	  8
	195 bp
	430 bp
	*12:01:01-12:01:03, 12:03:02-12:06, 12:08-12:11, 12:14, 12:17, 12:22, 12:24N-12:25
	*08:03:02, 08:10, 08:12, 08:14-08:15, 08:18-08:19, 08:23, 08:25, 08:27, 08:29-08:30:02, 08:32-08:38, 08:40, 08:45

	
	
	
	
	

	  9
	165 bp
	430 bp
	*12:01:01, 12:02:01, 12:02:03-12:02:05, 12:04, 12:06-12:13, 12:17-12:21, 12:23-12:27
	

	10
	195 bp
	430 bp
	*12:02:01-12:02:05, 12:13, 12:15-12:16, 12:18-12:21, 12:23, 12:26-12:27
	*08:01:01-08:02:03, 08:04:01-08:09, 08:11, 08:16-08:17, 08:21-08:22, 08:24, 08:26, 08:28, 08:31, 08:39, 08:41-08:44, 11:67, 14:15, 14:73

	11
	250 bp
	430 bp
	*12:03:02, 12:19
	*08:04:01, 08:04:02w-08:04:03w, 08:04:04-08:04:07, 08:06, 08:10, 08:28, 08:31, 11:67, 13:17, 13:116, 14:04, 14:11, 14:15, 14:31, 14:50, 14:52, 14:73, 14:76, 14:79, 14:107

	126
	170 bp
	430 bp
	*12:04
	*08:31, 08:41, 11:67, 14:11

	135,9
	115 bp, 185 bp, 255 bp
	515 bp
	*12:05, 12:14-12:15, 12:20, 12:24N
	

	146,8
	135 bp
	515 bp
	*12:06
	*15:01:01:01w-15:01:16w, 15:02:01w-15:57w, 16:01:01w-16:05:02w, 16:07w-16:18w

	15
	200 bp
	515 bp
	*12:07
	

	165,7
	75 bp
	430 bp
	*12:08, 12:23, 12:27
	*11:76, 13:34, 13:64, 14:41, 14:77, 14:110

	
	
	
	
	

	175,7
	90 bp
	430 bp
	*12:10, 12:25
	

	185,10
	90 bp, 

135 bp
	515 bp
	*12:11, 12:26
	

	19
	195 bp
	430 bp
	*12:12
	*08:13

	20
	220 bp
	430 bp
	*12:13, 12:23
	

	215
	105 bp
	430 bp
	*12:01:01-12:13, 12:15-12:27
	*08:32

	225,11
	105 bp, 220 bp
	430 bp
	*12:16, 12:21-12:22
	*08:32

	235
	110 bp
	430 bp
	*12:17, 12:25
	*01:01:01-01:36

	24
	170 bp
	430 bp
	*12:18
	


1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of DRB1*12 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective lengths of the specific PCR product(s) of the alleles amplified by these primer mixes are given. 

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 430 base pairs, for most wells, or a band of 515 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the longer, 515 bp, internal positive control band in order to help in the correct orientation of the DRB1*12 subtyping.

In addition, wells number 5, 13 to 15 and 18 contain the primer pair giving rise to the longer, 515 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3For several DRB alleles only partial second exon and third exon nucleotide sequences are available. In these instances it is not known whether some of the primers of the SSP sets are completely matched with the target sequences or not. We assume that unknown sequences in the first hyperpolymorphic region of the second exon of DRB alleles are conserved within allelic groups and that unknown sequences of codons 87 to 92 are identical with the DRB1*01:01 consensus sequence. 

4Due to the sharing of sequence motifs, non-DRB1*12 alleles will be amplified by primer mixes 1 to 6, 8, 10 to 12, 16, 19 and 21 to 23. In addition, the DRB1*15 and DRB1*16 alleles will be weakly amplified by primer mix 14.

5Short specific PCR fragments have a lower intensity than longer PCR bands.

6Primer mixes 12 and 14 may give rise to a primer dimer artefact.
7Primer mixes 16 and 17 may have tendencies of giving rise to unspecific amplifications.
8Primer mix 14 may give a lower yield of HLA-specific PCR product than the other DRB1*12 primer mixes.
9Primer mix 13: Specific PCR fragment of 115 bp in the DRB1*12:24N allele. Specific PCR fragment of 185 bp in the DRB1*12:05, 12:14 and 12:15 alleles. Specific PCR fragment of 255 bp in the DRB1*12:20 allele.

10Primer mix 18: Specific PCR fragment of 90 bp in the DRB1*12:26 allele. Specific PCR fragment of 135 bp in the DRB1*12:11 allele.

11Primer mix 22: Specific PCR fragment of 105 bp in the DRB1*12:21 allele. Specific PCR fragment of 220 bp in the DRB1*12:16 and the DRB1*12:22 alleles. Specific PCR fragment of 105 and 220 bp in the DRB1*08:32 allele.

 ‘w’, may be weakly amplified.

[image: image1.emf]Well No. 1 2 3 4 5 6 7 8 9 101112131415161718192021222324

DRB1 allele

1 

*12:01:01 1 2 4 5 6 7 8 9 21

*12:01:02-12:01:03 1 2 4 5 6 7 8 21

*12:02:01, 12:02:03-12:02:05 1 2 4 5 6 7 9 10 21

*12:02:02 1 2 4 5 6 7 10 21

*12:03:02 1 2 4 5 8 11 21

*12:04 1 2 4 6 7 8 9 12 21

*12:05 1 2 5 6 7 8 13 21

*12:06 1 2 4 5 6 7 8 9 14 21

*12:07 1 2 4 5 6 7 9 15 21

*12:08 1 2 4 5 6 8 9 16 21

*12:09 1 2 3 4 6 7 8 9 21

*12:10 1 2 4 5 6 7 8 9 17 21

*12:11 1 2 4 5 6 7 8 9 18 21

*12:12 1 2 4 5 6 7 9 19 21

*12:13 1 2 4 5 6 w 9 10 2021

*12:14 1 2 5 6 7 8 13

*12:15 1 2 5 6 7 10 13 21

*12:16 1 2 4 5 10 2122

*12:17 1 2 4 5 6 7 8 9 21 23

*12:18 1 2 4 6 7 9 10 21 24

*12:19 1 2 4 5 9 1011 21

*12:20 1 2 4 5 6 7 9 10 13 21

*12:21 2 4 5 6 7 9 10 2122

*12:22 1 4 5 8 2122

*12:23 1 2 4 5 6 9 10 16 2021

*12:24N 1 2 4 w 6 7 8 9 13 21

*12:25 1 2 4 5 6 7 8 9 17 21 23

*12:26 1 2 4 5 6 7 9 10 18 21

*12:27 1 2 4 5 6 9 10 16 21

*01:01:01-01:36 23

*08:01:01-08:01:05, 08:07-08:08, 

08:11, 08:16, 08:26, 08:39, 08:43

10

*08:02:01-08:02:03, 08:09, 08:21, 

08:42, 08:44

3 10

*08:03:02, 08:14-08:15, 08:23, 

08:27, 08:29, 08:30:02, 08:33, 

08:35-08:36, 08:38

8

*08:04:01, 08:04:04-08:04:05, 

08:04:07, 14:15

3 1011

*08:04:02-08:04:03 3 10w

*08:04:06, 08:06, 14:73 1011

*08:05 2 10

*08:10 8 11

*08:12 6 8

*08:13 3 19

*08:17 1 10

*08:18, 08:40 2 8

*08:19, 08:34 5 8

*08:22 6 10

*08:24 2 3 10

*08:25 2 5 8

*08:28 1 3 1011

*08:30:01 3 8

*08:31 2 101112

Well No. 1 2 3 4 5 6 7 8 9 101112131415161718192021222324



[image: image2.emf]Well No. 1 2 3 4 5 6 7 8 9 101112131415161718192021222324

*08:32 4 8 2122

*08:37 1 8

*08:41 2 10 12

*08:45 1 3 8

*11:67 1 101112

*11:76, 13:34, 13:64, 14:41, 14:77, 

14:110

16

*13:17 1 2 3 11

*13:116, 14:52 2 3 11

*14:04, 14:50, 14:76, 14:79, 

14:107

11

*14:11 1112

*14:28 6

*14:31 2 11

*15:01:01:01-15:01:16, 15:02:01-

15:57, 16:01:01-16:05:02, 16:07-

16:18

w

Length of spec. PCR product(s)

2

135215165105165250215195165195250170115135200

759090

195220105105110170

185 135 220

255

Well No. 1 2 3 4 5 6 7 8 9 101112131415161718192021222324

Specific Reaction

1st match: DRB1*

2nd match: DRB1*

Control Reaction

Failed Controls

False Positive

False Negative


1The sequence of the DRB1*12031 allele has been shown to be identical to DRB1*12:01:01.
2Primer mix 13: Specific PCR fragment of 115 bp in the DRB1*12:24N allele. Specific PCR fragment of 185 bp in the DRB1*12:05, 12:14 and 12:15 alleles. Specific PCR fragment of 255 bp in the DRB1*12:20 allele.

Primer mix 18: Specific PCR fragment of 90 bp in the DRB1*12:26 allele. Specific PCR fragment of 135 bp in the DRB1*12:11 allele.

Primer mix 22: Specific PCR fragment of 105 bp in the DRB1*12:21 allele. Specific PCR fragment of 220 bp in the DRB1*12:16 and the DRB1*12:22 alleles. Specific PCR fragment of 105 and 220 bp in the DRB1*08:32 allele.

 ‘w’, may be weakly amplified.






Olerup SSP AB

For In Vitro Diagnostic Use 
Franzengatan 5

www.olerup-ssp.com
SE-112 51 Stockholm
Sweden

