Gel Documentation Form and Worksheet

DRB3 (101.121-24/04,-24u/04u)	Lot No: 7F6	Expiry Date: 2020-06-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 9, 10 and 13 may have tendencies of unspecific amplifications.

Primer mix 7, 11, 19, 22 and 46 have a tendency giving rise to primer oligomer formation.

Primer mixes 4, 10 and 29 may give rise to a lower yield of HLA-specific PCR product than the other DRB3 primer mixes.

Primer mix 46 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1DRB3 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.29.0, August 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following DRB3 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified DRB3 alleles | Other amplified DRB alleles |
| --- | --- | --- | --- |
| 5 | 95 bp 125 bp | *01:03 *01:10 | DRB1*14:46 |
| 7 | 90 bp 180 bp | *01:05 *01:07, 01:15, 02:42 | DRB1*03:42, DRB1*14:46 |
| 12 | 90 bp 145 bp | *02:17 *02:01, 02:04, 02:24, 02:48, 02:58 | DRB1*03:37, DRB1*03:80, DRB1*03:140, DRB1*04:183, DRB1*11:43, DRB1*11:50, DRB1*11:115, DRB1*11:151, DRB1*11:171, DRB1*11:174, DRB1*12:34, DRB1*13:131, DRB1*13:150, DRB1*14:59, DRB1*14:86, DRB1*14:96, DRB1*14:145, DRB1*14:153-14:154, DRB1*15:65 |
| 14 | 185 bp 265 bp | *01:07, 01:09, 02:02:01:01-02:02:09, 02:05, 02:07-02:13, 02:15-02:19, 02:21-02:23, 02:25-02:36, 02:39-02:47, 02:49, 02:53-02:57, 02:60, 02:62-02:65 *01:02 | DRB1*10:01:01:01w-10:01:09w, DRB1*10:03w-10:11w, DRB1*10:13w, DRB1*10:15w-10:16w, DRB1*10:18w-10:22w, DRB1*13:216, DRB1*14:141, DRB1*16:30 |
| 15 | 120 bp 180 bp 210 bp | *02:03, 2:38, 02:50, 03:01:01-03:09, 03:11, 03:13 *02:16 *02:17 | DRB1*12:57, DRB1*13:195 |
| 17 | 120 bp 180 bp | *01:08, 02:06, 02:20, 03:10 *02:16, 02:23 | DRB1*03:42, DRB1*03:87, DRB1*13:67, DRB1*14:46 |
| 18 | 100 bp 170 bp | *01:09, 02:07, 02:09, 02:21, 02:44, 02:56 *01:18 |  |
| 19 | 180 bp 270 bp | *01:42, 02:08, 02:18, 02:23 *02:14 | DRB1*11:30 |
| 21 | 120 bp 155 bp 210 bp | *01:13 *02:43 *01:19, 01:31, 02:11 |  |
| 22 | 185 bp 245 bp | *01:11 *01:02, 02:12 |  |
| 23 | 195 bp 240 bp | *02:13 *01:02 |  |
| 25 | 150 bp 220 bp 260 bp | *02:60 *01:19, 02:26 *01:12, 02:31:01 |  |
| 28 | 235 bp 260 bp | *02:27 *01:12, 02:31:01, 02:55N | DRB1*11:30, DRB1*12:57, DRB1*13:67, DRB1*13:195, DRB1*14:46 |
| 31 | 70 bp 95 bp | *01:21, 02:19, 02:21, 02:25, 02:48, 02:56 *01:16, 02:29N | DRB1*12:34 DRB1*01:77w |
| 33 | 150 bp 180 bp | *01:32, 01:39, 01:41, 02:37-02:38, 02:50, 03:01:01-03:02, 03:05-03:09, 03:11-03:13 *01:20 | DRB1*03:35, DRB1*07:04, DRB1*14:38:01-14:38:02, DRB1*14:50, DRB1*14:93, DRB1*14:127:01-14:127:02 |
| 34 | 135 bp 175 bp | *01:43, 03:05 *01:24, 02:36 |  |
| 35 | 115 bp 250 bp | *02:51 *01:17, 02:33 |  |
| 36 | 105 bp 155 bp 185 bp | *01:22 *02:43 *01:35, 02:35 |  |
| 37 | 155 bp 215 bp | *02:30 *01:31, 01:45, 02:11, 02:47 |  |
| 38 | 70 bp 205 bp | *02:32 *01:07?, 02:05?, 02:07?-02:08?, 02:34 | DRB1*03:44, DRB1*10:13, DRB1*13:106 |
| 39 | 190 bp 260 bp | *02:39 *01:40N |  |
| 40 | 130 bp 190 bp | *03:13 *01:07, 01:15, 01:30, 01:42, 02:42 | DRB1*03:42, DRB1*14:46 |
| 42 | 165 bp 205 bp | *02:61Q *01:37, 03:08 |  |
| 44 | 100 bp 260 bp | *01:26N *02:55N |  |

4The following DRB3 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| DRB3*02:32, 02:34 | 38 |

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.