Gel Documentation Form and Worksheet

DRB5 (101.123-24/06, -24u/06u) 	Lot No: 5F2	Expiry Date: 2020-04-01

Name:___________________

Sample ID:_______________

DNA Conc.(ng/ul):_________

Test Date: ______________

Tested By: ______________

Review Date:________________

Reviewed By:________________

Interpretation:__________	                       Failed lanes: ___________                             Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 19 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1DRB5 alleles listed on the IMGT/HLA web page 2017-April-13, release 3.28.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following DRB5 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified DRB5 alleles | Other amplified DRB alleles |
| --- | --- | --- | --- |
| 4 | 100 bp 150 bp | *01:01:01-01:01:03, 01:04, 01:06-01:07, 01:09, 01:11, 01:15-01:19, 01:21-01:24, 01:26, 01:29-01:31 *02:06 |  |
| 9 | 85 bp 175 bp 225 bp | *01:16 *01:13 *01:04 |  |
| 10 | 130 bp 160 bp | *01:07 *01:12, 01:15 |  |
| 11 | 110 bp 200 bp | *01:14 *01:06, 01:11, 01:21, 02:02-02:03, 02:06-02:07, 02:09-02:11, 02:13-02:16 |  |
| 13 | 150 bp 195 bp | *01:01:02?, 01:03?, 01:07?, 01:09?, 01:18, 02:04? *01:08N | DRB1*15:02:03?, DRB1*15:86, DRB1*16:01:02?, DRB1*16:02:02?, DRB1*16:05:01?, DRB4*01:05?, DRB1*01:07:01? |
| 16 | 125 bp 225 bp | *01:21, 02:10 *01:20, 02:08 | DRB1*08:59, DRB1*11:210 |
| 17 | 130 bp 180 bp | *02:12 *01:27N |  |

‘?’, nucleotide sequence information not available for the primer matching sequence.