# HLA-A low resolution 	Lot No: 9R2	Expiry Date: 2027-03-01

# (101.401-48/12,-48u/12u)

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

Interpretation:__________	Failed lanes: ___________ 	   Comments:_______________

Gel Picture

## PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 6 may give rise to a lower yield of A*23 alleles than the other A low primer mixes.

Primer mix 9 may weakly amplify the A*34 alleles.

Primer mix 22 might faintly amplify most A*11 alleles.

Primer mixes 13, 15, 24, 27 and 28 may give rise to a lower yield of HLA-specific PCR product than the other HLA-A low primer mixes.

Primer mixes 6, 10, 14, 20, 23 and 30 have a tendency to giving rise to primer oligomer formation.

Primer mixes 15, 17, 20, 24 and 27 to 29 may have tendencies of unspecific amplifications.

Primer mix 30 may generate a false positive band of about 500 base pairs. This band should be disregarded when interpreting HLA-A low resolution typings.

Primer mix 32 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR products generated by the HGH positive control primer pair is 200 base pairs.

1HLA-A alleles listed on the IMGT/HLA web page 2023-January-12, release 3.51.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all HLA-A alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele.

4The following alleles give rise to identical amplification patterns with the HLA-A low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

| Alleles | Alleles |
| --- | --- |
| A*01:427, 36:02 | A*25:78, 26:57, 66:15 |
| A*11:325, 24:555 | A*30:12:01-30:12:02, 30:31, 30:152, 30:166, B*07:260, C*12:328 |
| A*23:14:01-23:14:02, 23:104-23:105, 24:71, 24:315, 24:392, 24:527 | A*31:08, 33:53 |
| A*23:66, 23:99, 23:128, 24:14:01:01-24:15, 24:51-24:53, 24:57, 24:64, 24:94, 24:114, 24:138, 24:188, 24:222N, 24:228, 24:291, 24:296, 24:304, 24:316, 24:324, 24:412, 24:481, 24:515, B*53:72, C*04:01:03 | A*34:06, 66:40 |

Abbreviations

ser: serological HLA specificity

Changes in revision R01 compared to R00:

Primer mix 27 does not amplify the A*03:69N allele. This has been corrected in the Product Insert and the Worksheet