# HLA-A low resolution 	Lot No: 0V4	Expiry Date: 2028-11-01

# (101.401-48/12,-48u/12u)

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

Interpretation:__________	Failed lanes: ___________ 	   Comments:_______________

Gel Picture

## PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 6 may give rise to a lower yield of A*23 alleles than the other A low primer mixes.

Primer mix 9 may weakly amplify the A*34 alleles.

Primer mix 22 might faintly amplify most A*11 alleles.

Primer mixes 15, 24, 27 and 28 may give rise to a lower yield of HLA-specific PCR product than the other HLA-A low primer mixes.

Primer mixes 6, 10, 11, 14, 20, 23 and 30 have a tendency to giving rise to primer oligomer formation.

Primer mixes 13, 15, 17, 20, 23, 24 and 27 to 30 may have tendencies of unspecific amplifications, most pronounced in primer mix 24.

Primer mix 30 may generate a false positive band of about 500 base pairs. This band should be disregarded when interpreting HLA-A low resolution typings.

Primer mix 32 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1HLA-A alleles listed on the IMGT/HLA web page 2024-April-10, release 3.56.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The serological reactivity of all HLA-A alleles is not known. In this table we use the expert-assigned serological grouping in Tissue Antigens (2009) 73:95-170 and the serological grouping of the sequence-defined allele.

4The following alleles give rise to identical amplification patterns with the HLA-A low resolution primer set. These alleles can be separated by the respective high resolution primer sets.

| Alleles | Alleles |
| --- | --- |
| A*01:427, 36:02 | A*30:12:01-30:12:02, 30:31, 30:152, 30:166, B*07:260, C*12:328 |
| A*11:325, 24:555 | A*31:08, A*33:53 |
| A*23:14:01-23:14:02, 23:104-23:105, A*24:71, 24:315, 24:392, 24:527 | A*31:109, A*33:249 |
| A*23:66, 23:99, 23:128, 23:137, A*24:14:01:01-24:15, 24:51-24:53, 24:57, 24:64, 24:94, 24:114, 24:138, 24:188, 24:222N, 24:228, 24:291, 24:296, 24:304, 24:316, 24:324, 24:412, 24:481, 24:515, 24:610, B*53:72, C*04:01:03 | A*34:06, A*66:40 |
| A*25:78, A*26:57, A*66:15 |

Abbreviations

ser: serological HLA specificity

Changes in revision R01 compared to R00:

Primer mix 27 does not amplify the A*03:69N allele. This has been corrected in the Product Insert and the Worksheet