# Gel Documentation Form and Worksheet

# HLA-A*01 (101.411-24/06, -24u/06u) 	Lot No: 6F2	Expiry Date: 2022-04-01

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

# Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

## PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 10, and 51 have a tendency of giving rise to primer oligomer formation.

Primer mixes 5, 8, 9,16, 22, 28, 45 and 47 may have tendencies of unspecific amplifications.

Primer mix 12 may give rise to a lower yield of HLA-specific PCR product than the other A*01 primer mixes.

Primer mix 60 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

# Change in revision R01 compared to R00:

# Primer mix 30 does not amplify the A*01:37 and the A*11:172 alleles. This has been corrected in the Specificity and Interpretation Tables. Thus, this lot of the HLA-A*01 subtyping kit cannot distinguish the A*01:01:01:01, 01:01:01:03-01:01:22, 01:01:24-01:01:37, 01:01:39-01:01:47, 01:01:49-01:01:78, 01:01:80-01:01:82 alleles and the A*01:37 allele.

# Change in revision R02 compared to R01.

# Primer mix 1 amplifies the A*01:15N allele. This has been corrected in the Specificity and Interpretation Tables.

Changes in revision R03 compared to R02:

1.	The expiration date has been altered due to extension of shelf-life.

1HLA-A*01 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.24.0, May 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-A*01 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-A*01 alleles | Other amplified HLA-A alleles |
| --- | --- | --- | --- |
| 6 | 125 bp 210 bp | *01:60 *01:09:01-01:09:02 | *26:31, C*06:71, C*07:581 C*07:274 |
| 7 | 60 bp 115 bp | *01:06 *01:86 | *02:576, 31:62 |
| 8 | 110 bp 180 bp | *01:07, 01:23, 01:51, 01:83:01-01:83:02, 01:191 *01:01:01:02N | *24:243, 26:120, 31:35 |
| 10 | 120 bp 155 bp 240 bp 270 bp | *01:150 *01:10 *01:137 *01:29 | *03:231:01w -03:231:02, 11:14w, 11:50Q, 30:26, 80:01:01:01w-80:03w |
| 11 | 135 bp 180 bp 275 bp | *01:13, 01:28, 01:176, 01:194, 01:229 *01:106 *01:11N | *31:35 |
| 12 | 90 bp 125 bp | *01:86, 01:115 | *02:576, 02:682, 03:187, 11:155, 11:226, 31:62, 36:01-36:05, 68:41, B*40:359, B*57:65, C*04:31, C*06:137, C*07:569 B*40:359, C*07:569w |
| 14 | 75 bp 120 bp | *01:59 *01:13, 01:17, 01:176, 01:194 |  |
| 16 | 180 bp 235 bp | *01:01:01:02N *01:15N |  |
| 17 | 180 bp 210 bp 285 bp | *01:106 *01:16N *01:101 | *03:87, 11:30, 30:92 |
| 21 | 125 bp 220 bp 255 bp | *01:44 *01:155 *01:20, 01:66, 01:130 | *02:19, 02:36-02:37, 02:54, 02:255, 02:417, 24:14:01:01-24:14:01:03, 24:93, 24:324 |
| 25 | 80 bp 425 bp | *01:31N, 01:51, 01:59 *01:19, 01:173 | *26:120 |
| 26 | 90 bp 460 bp 545 bp | *01:104, 01:134, 01:229 *01:32 *01:45 | *02:346, 02:427, 11:06, 25:11, 26:03:01, 26:06, 26:21, 26:36, 26:78, 26:92, 26:111, 26:146, 80:01:01:01w *02:453, 02:557, 02:690, 03:78, 11:108, 24:271, 66:17 |
| 27 | 110 bp 170 bp 195 bp | *01:87N *01:33, 01:141 *01:109 | *02:679, 03:193 *03:182, 11:100, 11:175 |
| 28 | 95 bp 135 bp 195 bp | *01:01:38L *01:132 *01:109 | *24:02:03Q *03:182, 11:100, 11:175 |
| 29 | 110 bp 155 bp 230 bp 285 bp | *01:69:01-01:69:02 *01:35 *01:68 *01:101 | *02:609, 03:107, 11:17, 23:09, 24:129, 26:62, 26:72, 32:24 *03:77, 11:144 *03:87, 11:30, 30:92 |
| 30 | 135 bp 175 bp 220 bp 280 bp | *01:95 *01:07, 01:100, 01:190, 01:194 *01:155 *01:142 | *24:243 *24:243, 31:35 *24:150 |
| 32 | 110 bp 140 bp | *01:57N *01:43 | B*40:291N, C*06:152N, C*07:191N *26:120 |
| 34 | 130 bp 215 bp | *01:60, 01:71, 01:115 *01:58 | *26:31, B*40:359, C*06:71, C*07:569w, C*07:581 |
| 35 | 110 bp 180 bp 240 bp | *01:47, 01:150 *01:07, 01:49, 01:100, 01:190, 01:194 *01:137 | *11:229, 24:243, 31:35 |
| 36 | 90 bp 140 bp 175 bp 235 bp | *01:54 *01:48 *01:53N *01:123N | *03:250 |
| 37 | 65 bp 120 bp 155 bp | *01:65 *01:71 *01:77, 01:92 |  |
| 38 | 100 bp 180 bp 240 bp | *01:50 *01:62 *01:68, 01:72 | *02:315, 03:85 *11:236, 30:45 |
| 39 | 125 bp 170 bp | *01:77 *01:39, 01:113 | B*07:241, C*04:251 *24:26, 24:314 |
| 42 | 110 bp 185 bp 255 bp | *01:87N *01:64, 01:141 *01:142 | *03:232 *24:150 |
| 43 | 550 bp 590 bp | *01:45, 01:56N *01:107 | *02:453, 02:557, 02:690, 03:78, 11:69N, 11:108, 24:271, 66:17 |
| 44 | 65 bp 245 bp | *01:98 *01:46 | *24:87, 24:285, B*07:64 |
| 45 | 65 bp 140 bp | *01:98 *01:52:01N | *24:87, 24:285, B*07:64 |
| 46 | 85 bp 155 bp 245 bp | *01:81 *01:40 *01:72, 01:147Q | *03:26, 11:77, 11:126 *11:160 *11:236, 30:45 |
| 47 | 295 bp 325 bp 450 bp 550 bp | *01:41 *01:42 *01:48 *01:123N | *11:48 |
| 54 | 70 bp 155 bp | *01:103 *01:132 |  |

4The following HLA-A*01 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| A*01:31N, 01:173 | 25 | A*01:53N-01:54 | 36 |
| A*01:32, 01:104 | 26 | A*01:65, 01:92 | 37 |
| A*01:47, 01:49 | 35 | A*01:87N, 01:141 | 27, 42 |
| A*01:48, 01:123N | 36 | A*01:137, 01:150 | 10, 35 |
| A*01:50, 01:62 | 38 |

3This lot of the HLA-A*01 subtyping kit cannot distinguish the A*01:01:01:01, 01:01:01:03-01:01:22, 01:01:24-01:01:37, 01:01:39-01:01:47, 01:01:49-01:01:78, 01:01:80-01:01:82 alleles and the A*01:37 allele.

‘w’, might be weakly amplified.