# Gel Documentation Form and Worksheet

# HLA-A*01 (101.411-24/06, -24u/06u) 	Lot No: 6H3	Expiry Date: 2021-11-01

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

# Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

## PHOTO DOCUMENT

Abbreviations

‘ICB’ Internal Control Band

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

# HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

# Primer mixes 10, 31 and 51 have a tendency of giving rise to primer oligomer formation.

# Primer mixes 5, 8, 9, 16, 22, 28, 45 and 47 may have tendencies of unspecific amplifications.

# Primer mix 5, 12 and 13 may give rise to a lower yield of HLA-specific PCR product than the other A*01 primer mixes.

# Primer mix 62 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-A*01 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.24.0, May 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-A*01 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-A*01 alleles | Other amplified HLA-A alleles |
| --- | --- | --- | --- |
| 6 | 125 bp 210 bp | *01:60 *01:09:01:01-01:09:02 | *26:31, C*06:71, C*07:581 C*07:274 |
| 7 | 60 bp 115 bp | *01:06 *01:86 | *02:576, 31:62 |
| 8 | 110 bp 180 bp | *01:07, 01:23, 01:51, 01:83:01-01:83:02, 01:191 *01:01:01:02N | *24:243, 26:120, 31:35 |
| 10 | 120 bp 155 bp 240 bp 270 bp | *01:150 *01:10 *01:137 *01:29 | *03:231:01w, 03:231:02, 11:14w, 11:50Q, 11:290, 30:26, 80:01:01:01w-80:04w |
| 11 | 135 bp 180 bp 275 bp | *01:13, 01:28, 01:176, 01:194, 01:229 *01:106 *01:11N | *11:282, 31:35 |
| 12 | 90 bp 125 bp | *01:86, 01:115 | *02:576, 02:682, 03:187, 11:155, 11:226, 31:62, 36:01-36:08, 68:41, B*40:359, B*57:65, C*04:31, C*06:137, C*07:569 B*40:359, B*44:289, C*07:569w |
| 14 | 75 bp 120 bp | *01:59 *01:13, 01:17, 01:176, 01:194 | *11:282 |
| 16 | 180 bp 235 bp | *01:01:01:02N *01:15N |  |
| 17 | 180 bp 210 bp 285 bp | *01:106 *01:16N *01:101 | *03:87, 11:30, 30:92 |
| 19 | 150 bp 170 bp | *01:269N *01:18N, 01:30, 01:145, 01:236 | *26:145N *02:262, 02:547, 02:700, 24:22, 24:160, 24:299, 26:149 |
| 21 | 125 bp 220 bp 255 bp | *01:44 *01:155 *01:20, 01:66, 01:130 | *02:19, 02:36-02:37, 02:54, 02:255, 02:417, 02:723, 24:14:01:01-24:14:01:03, 24:93, 24:324 |
| 25 | 80 bp 425 bp | *01:31N, 01:51, 01:59 *01:19, 01:173 | *26:120 |
| 26 | 90 bp 460 bp 545 bp | *01:104, 01:134, 01:229 *01:32 *01:45 | *02:346, 02:427, 11:06, 25:11, 26:03:01, 26:06, 26:21, 26:36, 26:78, 26:92, 26:111, 26:146, 80:01:01:01w *02:453, 02:557, 02:690, 03:78, 11:108, 23:92, 24:271, 33:151, 66:17 |
| 27 | 110 bp 170 bp 195 bp | *01:87N *01:33, 01:141 *01:109 | *02:679, 03:193 *03:182, 11:100, 11:175 |
| 28 | 95 bp 135 bp 195 bp | *01:01:38L *01:132 *01:109 | *24:02:03Q *30:142 *03:182, 11:100, 11:175 |
| 29 | 110 bp 155 bp 230 bp 285 bp | *01:69:01-01:69:02 *01:35 *01:68 *01:101 | *02:609, 03:107, 11:17, 23:09, 24:129, 26:62, 26:72, 32:24 *03:77, 11:144 *03:87, 11:30, 30:92 |
| 30 | 135 bp 180 bp 220 bp 280 bp | *01:95 *01:07, 01:100, 01:190, 01:194 *01:155 *01:142 | *24:243 *24:243, 31:35 *24:150 |
| 32 | 110 bp 140 bp | *01:57N *01:43 | B*40:118N, B*40:291N, C*06:152N, C*07:191N *11:282, 26:120 |
| 34 | 130 bp 150 bp 215 bp | *01:60, 01:71, 01:115 *01:269N *01:58 | *26:31, B*40:359, B*44:289, C*06:71, C*07:569w, C*07:581 *26:145N |
| 35 | 110 bp 180 bp 240 bp | *01:47, 01:150 *01:07, 01:49, 01:100, 01:190, 01:194 *01:137 | *11:229, 24:243, 31:35 |
| 36 | 90 bp 140 bp 175 bp 235 bp | *01:54 *01:48 *01:53N *01:123N | *03:250 |
| 37 | 65 bp 115 bp 150 bp | *01:65 *01:71, 01:247N *01:77, 01:92 |  |
| 38 | 100 bp 180 bp 240 bp | *01:50 *01:62 *01:68, 01:72 | *02:315, 03:85 *11:236, 30:45 |
| 39 | 125 bp 170 bp | *01:77 *01:39, 01:113 | B*07:241, C*04:251 |
| 42 | 110 bp 185 bp 255 bp | *01:87N *01:64, 01:141 *01:142 | *03:232 *24:150 |
| 43 | 550 bp 590 bp | *01:45, 01:56N *01:107 | *02:453, 02:557, 02:690, 03:78, 11:69N, 11:108, 23:92, 24:271, 33:151, 66:17 |
| 44 | 65 bp 245 bp | *01:98 *01:46 | *24:87, 24:285, B*07:64 |
| 45 | 65 bp 140 bp | *01:98 *01:52:01N | *24:87, 24:285, B*07:64 |
| 46 | 85 bp 155 bp 245 bp | *01:81 *01:40 *01:72, 01:147Q | *03:26, 03:318, 11:77, 11:126 *11:160 *11:236, 30:45 |
| 47 | 295 bp 325 bp 450 bp 550 bp | *01:41 *01:42 *01:48 *01:123N | *11:48 |
| 50 | 200 bp 230 bp | *01:250N *01:162N |  |
| 53 | 110 bp 215 bp | *01:247N *01:186N |  |
| 54 | 70 bp 155 bp | *01:103 *01:132 | *30:142 |
| 55 | 150 bp 190 bp | *01:52:02N *01:258N |  |
| 57 | 190 bp 240 bp | *01:258N *01:208Q |  |
| 59 | 180 bp 200 bp | *01:240N *01:250N |  |

4The following HLA-A*01 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| A*01:31N, 01:173 | 25 | A*01:53N-01:54 | 36 |
| A*01:32, 01:104 | 26 | A*01:65, 01:92 | 37 |
| A*01:47, 01:49 | 35 | A*01:87N, 01:141 | 27, 42 |
| A*01:48, 01:123N | 36 | A*01:137, 01:150 | 10, 35 |
| A*01:50, 01:62 | 38 |

Abbreviations

‘w’, might be weakly amplified.

Changes in revision R01 compared to R00:

1. The well assignment of the following alleles was corrected in the interpretation table: A*01:35, 01:37, 01:228Q, 01:240N, 01:247N, 01:248Q, 01:250N, A*11:155, and 11:172.