PAGE  
[image: image1.emf]Well No. 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24

ICB 800 107010701070 800 10701070 800 800 1070 800 1070 800 1070 800 107010701070 800 800 10701070 800 1070

195 270 190 235 85 190 210 100 85 120 215 235 195 100 100 95 205 225 100 180 100 150 80 295

180 550 175 280 215 210 240 175 235 230 465

210 265

Well No. 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24

Specific 

Reaction

AmpS

[image: image2.emf]Well No. 25 26 27 28 29 30 31 32 33 34 35 36 37 38 39 40 41 42 43 44 45 46 47 48

ICB 107010701070 800 107010701070107010701070107010701070107010701070107010701070 800 1070107010701070

190 305 160 240 80 180 230 205 110 85 105 65 105 95 205 225 90 95 445 145 210 75 55 195

220 370 205 280 200 285 295 280 170 290 230 185 120 280 215 175 215 340 120

260 235 305 200 380 290 225 560 400 175

270 265 220

Well No. 25 26 27 28 29 30 31 32 33 34 35 36 37 38 39 40 41 42 43 44 45 46 47 48

Specific 

Reaction

AmpS

[image: image3.emf]Well No. 49 50 51 52 53 54 55 56 57 58 59 60 61 62 63 64

ICB 800 800 107010701070107010701070 800 107010701070107010701070

180 90 265 230 150 200 115 65 75 95 165 105 130 205 145

230 210 545 195 270 145 230 175 245 340 270 265 170

310 170 245 370 275

400 235

Well No. 49 50 51 52 53 54 55 56 57 58 59 60 61 62 63 64

Specific 

Reaction

AmpS

NC

 ®





HLA-A*11
January 2017
101.416-12/04 -12u/04u

Rev. No: 02                                                                 62Y        


 Gel Documentation Form and Worksheet 
HLA-A*11 (101.416-12/04, -12u/04u) 
Lot No: 62Y
Expiry Date: 2018-January-01
Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________
 

Interpretation:___________                          Failed lanes: ____________                         Comments:________________       
Gel Picture
	PHOTO DOCUMENT


[image: image4.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*11:182Q 1 11 14 30 49

*11:183 1 7 12 14 32 41

*11:189 1 1112 14 38 42

*11:191 1 3 7 12 14 62

*11:194 1 1112 14 55 63

*11:199 1 1112 14 40

*11:201 2 8 1112 14

*11:208N 1 1112 14 27 34

*11:209 1 4 11 14 27 41

*11:210N 1 1112 14 36 45 50

*11:211 1 7 12 14 56

*11:215N 1 1112 14 19 30

*01:01:01:01-01:01:22, 

01:01:24-01:01:27, 01:01:29-

01:01:47, 01:01:49-01:01:56, 

01:01:59-01:01:61, 01:01:63-

01:02, 01:09, 01:11N, 01:16N, 

01:18N, 01:22N-01:24, 01:27N-

01:33, 01:35-01:36, 01:38-

01:40, 01:42, 01:44, 01:46-

01:50, 01:52N-01:55, 01:57N-

01:60, 01:62, 01:64-01:65, 

01:67:01-01:70, 01:72, 01:74-

01:78, 01:80-01:82, 01:84-

01:85, 01:87N-01:88, 01:90-

01:99, 01:102-01:103, 01:106-

01:108, 01:110-01:114, 01:116-

01:124, 01:128, 01:131-01:133, 

01:135, 01:137-01:142, 01:145-

01:146, 01:150-01:158, 

01:160N-01:175, 01:179N, 

36:04

7 12 21 27

*01:01:23, 01:01:48, 01:08, 

01:15N, 01:147Q, 01:149, 

01:159, 01:178N

7 21 27

*01:01:28, 01:01:57, 01:01:62, 

01:07, 01:71, 01:100, 01:105, 

01:144

7 12 27

*01:01:58, 01:20, 01:63, 01:66, 

01:73, 01:89, 01:130, 02:346, 

02:427

21

*01:03, 32:62, 36:03, 74:10 21 26

*01:04N 7 12 21 24 27

*01:06 12 15 21 27

*01:10 7 12 17 21 27 56

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control


[image: image5.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*02:05:01-02:06:04, 02:06:06-

02:06:18, 02:08, 02:10, 02:14, 

02:21, 02:28, 02:41, 02:44, 

02:51, 02:54, 02:57, 02:61, 

02:72, 02:79:01-02:79:02, 

02:84, 02:91, 02:106, 02:108, 

02:122, 02:126-02:127, 02:137, 

02:142-02:144, 02:169-02:170, 

02:172, 02:178-02:180, 02:229, 

02:232, 02:244, 02:248, 02:259, 

02:271, 02:278, 02:286, 02:290, 

02:295, 02:300, 02:310, 02:324, 

02:328, 02:330, 02:333, 02:337, 

02:344, 02:358-02:359, 

02:373N, 02:376, 02:382, 

02:387, 02:398, 02:404-02:405, 

02:409, 02:413, 02:415, 02:419-

02:421, 02:428, 02:433, 02:438, 

02:454, 02:465, 02:470-

02:476N, 02:493, 02:495-

02:496, 02:506N-02:507, 

02:532, 02:546, 02:550, 02:558, 

02:572, B*27:87

22

*02:26, 02:171:02 23

*02:34-02:35:03 1 w

*02:38, 02:101:01, 02:543 6 16

*02:55, 33:01:01-33:01:07, 

33:03:01-33:07, 33:10-33:12, 

33:14-33:16, 33:18:01-33:20, 

33:22-33:37, 33:39-33:47, 

33:50-33:52, 33:54-33:60, 

33:62-33:65, 33:67-33:69, 

33:71-33:97, 66:03, 68:03:01-

68:05, 68:20, 68:31, 68:90

29

*02:56:01-02:56:02, 02:62, 

02:78, 02:103, 31:89, 68:01:27, 

68:10, 68:13:01-68:14, 68:66

1 15

*02:65, 02:80, 02:114, 02:152, 

02:246, 02:279, 02:298, 02:304, 

02:406, 29:48, 33:08-33:09, 

74:04, 74:21, B*57:04, B*57:32

43

*02:99 2223

*02:117 w 43

*02:135 6

*02:154 6 16 22

*02:156, 02:338 4 15 32 37 41 49

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control


[image: image6.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*03:191 1 4 15 4041 62

*03:198 4 11 41 43

*03:208 1 15 40 55 62

*03:210 1 4 15 25 40 48 62

*03:218 1 4 11 15 4041 50

*23:07N, 24:11N 24

*23:10, 24:46 w

*24:10:01-24:10:02, 24:210, 

24:300, 31:24

16

*24:19, C*03:82 1

*24:24, 29:01:01:01-29:03, 

29:05-29:14, 29:16-29:18, 

29:21-29:29, 29:31-29:33, 

29:35-29:47, 29:49-29:65, 

29:67-29:73, 29:75-29:78N, 

30:52, 30:72, 32:17, 68:36, 

68:46, 68:84

15

*24:44, 24:109, 24:260 w

*24:59 9

*24:92 4 11 27 41

*24:177, 24:289, B*15:173, 

B*50:11, C*12:74

61

*24:249 55

*24:285 9 47

*24:290, 66:02, 66:12, 66:16, 

66:21, 68:01:01:01-68:01:17, 

68:01:19-68:01:26, 68:01:28, 

68:01:30-68:02:09, 68:06-

68:09, 68:11N-68:12, 68:16-

68:19, 68:21:01-68:23, 68:25-

68:29, 68:32-68:35, 68:37-

68:39, 68:41, 68:43-68:45, 

68:47-68:53, 68:55:01-68:56, 

68:58-68:60, 68:62, 68:64-

68:65, 68:67-68:70, 68:73-

68:74, 68:76:01-68:83, 68:86-

68:89, 68:91, 68:95-68:96, 

68:98-68:102, 68:104:01-

68:108, 68:110-68:122, 68:124-

68:126, 68:128-68:129, 68:131-

68:132, 69:01, 69:03

1 15 29

*25:01:01-25:01:09, 25:03-

25:09, 25:12N, 25:14, 25:16, 

25:18-25:31, 26:01:30, 26:05, 

26:16, 26:20, 26:57

10 21

*25:02 10 14

*25:10 10 16 21

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control


‘ICB’ Internal Control Band, 
‘AmpS’ Amplicon Size
Notes:
Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.
Primer mixes 9, 15, 16, 23, 36, 46, 53, 57, 62 and 63 may have tendencies of unspecific amplifications.
Primer mixes 7, 14, 21, 25 and 55 may give rise to a lower yield of HLA-specific PCR product than the other A*11 primer mixes.

Primer mixes 8, 9, 13, 16, 19, 27, 43, 50, 53, 55 and 60 may have a tendency of giving rise to primer oligomer formation. 

Primer mix 31 may give rise to a long unspecific amplification product of about 800 bp. This band should be disregarded when interpreting the A*11 SSP typings. 

Primer mix 64 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.
[image: image7.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*25:11 6 10 21

*25:15, 25:17, B*07:102, 

B*08:18, B*35:178, B*35:282, 

B*73:01-73:02, C*06:147

10

*26:01:01-26:01:21, 26:01:23-

26:01:27, 26:01:29, 26:01:31-

26:02:02, 26:04, 26:08-26:12, 

26:14-26:15, 26:17-26:18, 

26:22-26:27, 26:29, 26:34, 

26:37-26:43:02, 26:45-26:51, 

26:53-26:56, 26:58-26:61, 

26:63, 26:66-26:69, 26:73-

26:74, 26:79-26:82, 26:84-

26:91, 26:93-26:110

10 2021

*26:01:22, 26:01:28, 26:64, 

26:83

10 20

*26:03:01, 26:06, 26:21, 26:111

6 10 2122 29

*26:07:01-26:07:02, 26:31, 

26:76

2021

*26:13, 26:33 10 14 20

*26:19 14 20

*26:28, 26:52 10 2021 61

*26:30 6 22 29

*26:32, 26:70 10 2021 26

*26:35 10 21 28

*26:36 6 10 2021

*26:62 10 2021 27

*26:65 20

*26:71N 10 2021 27 34

*26:72 2021 27

*26:75 8 10 2021 50

*26:77 10 2021 38 42

*26:78 10 2122 29

*26:92 6 2122

*29:04 w

*29:15 8 15

*29:66 10 14 26 31

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



[image: image8.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*30:01:01-30:01:06, 30:01:08-

30:02:04, 30:02:06-30:03, 30:07-

30:08, 30:10-30:13, 30:15-

30:16, 30:19-30:20, 30:22, 

30:24-30:26, 30:30-30:33, 

30:36-30:38, 30:40-30:42, 

30:44-30:45, 30:48-30:50, 

30:53, 30:56-30:71, 30:73N-

30:74, 30:76N, 30:78N-30:79, 

30:81-30:85, 30:88, 30:91, 

30:93-30:94

3 4 15 17

*30:01:07, 30:23, 30:27N 3 4 15

*30:02:05 3 4 17 25

*30:04:01-30:04:02, 30:06, 

30:29, 30:46, 30:77, 30:90

3 4 7 32 w

*30:09, 31:04, 74:23 3 4 15 32 41

*30:14L, 30:43 3 15 17

*30:17 3 4 7 15 32 w

*30:18 1 3 4 15 17

*30:34 3 4 15 17 54 60

*30:35, 30:54, 30:86-30:87 4 15 17

*30:39 3 4 15 1718

*30:51 3 4 15 17 32

*30:55 1 3 4 15 17 40

*30:75 3 4 15 17 27

*30:92 3 4 15 17 31

*31:03 3 4 10 15 21 3132 41

*31:65, B*44:36 36

*33:13 10 21 26 31

*33:49 3 4 15 29 32 41

*33:61 16 29

*33:70 29 61

*34:01:01-34:01:02, 34:05, 

34:11

1 10 w 29

*34:02:01, 34:02:03-34:02:04, 

34:04, 34:10N

1 w 4 10 1415 29 62

*34:02:02 1 w 4 10 14 25 29 62

*34:03 1 3 4 10 1415 29 62

*34:06, 66:01:01-66:01:03, 

66:04, 66:06-66:07, 66:09-

66:10, 66:14, 66:18-66:20, 

66:22, 69:02

1 10 14 29

*34:07 1 w 4 10 15 29 62

*34:08 1 w 4 10 1415 29 40 62

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



[image: image9.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*34:09 1 w 4 15 29 62

*34:12 1 10 w 26 29

*43:01 13 2021

*66:05, 66:15 10 21 29

*66:08 1 5 10 14 29

*66:11 10 14 30

*66:13 1 10 14 29 38 58

*66:17 1 10 14 29 52

*68:01:18, 68:24, 68:40, 68:72, 

68:85, 68:93, 68:97, 68:123

1 29

*68:01:29, 68:57, 68:92 15 29

*68:15 22 29

*68:30, 68:75:01-68:75:02 1 15 43

*68:42, 68:54, 68:61 1 15 w 29

*68:63 1 15 23 29

*68:71 1 15 29 40

*68:94N 1 15 27 29 34

*68:103:01-68:103:02 1 4 7 15 29 32 37 w 49

*68:127 1 w 29

*68:130 1 13 15

*80:01:01:01-80:02 17 21 25 27 39 5455

*80:03 4 17 21 25 27 39 54

B*07:111N 19 30

B*07:117 38

B*15:90, B*45:05 w

B*39:75, B*40:268 1920

C*04:129, C*05:25, C*05:42, 

C*06:05, C*06:67, C*07:101, 

C*07:148, C*07:161, C*08:28

39 54

C*06:72, C*06:90, C*07:277, 

C*07:394

26

C*07:221 33

C*16:79 46

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



[image: image10.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*11:25 1 12 1415 23 62

*11:26 1 14 16 26

*11:27 1 4 9 11 14 27

*11:29 1 1112 14 30

*11:30 1 1112 14 31

*11:31 1 12 1415 32 41

*11:32 1 1112 14 32

*11:33:01-11:33:02 1 3 1112 14

*11:34 1 8 1112 14 50

*11:35 1 4 1415 32 41 49

*11:36 1 1112 14 57

*11:37 1 1112 14 22

*11:38 2 9 1112 14 27

*11:39 1 9 1112 14 27

*11:40 1112 2021

*11:41 1 1112 14 26

*11:42 1 1112 14 1920

*11:43 1 101112 14 31

*11:44 1 7 12 14 47

*11:45 1112 14 30

*11:46 1 1112 14 19 34

*11:47 1 1112 14 18 25

*11:48 1 14 24

*11:49 1 1112 14 28 31

*11:50Q 1 7 14 17 30 32 49

*11:51 1 1112 14 51

*11:52Q 1 1112 14 39

*11:54 1 1112 14 16 61

*11:55 1 1112 14 27 30

*11:56 1 1112 14 33

*11:57 2 1112 14 37 49

*11:58, 11:674 1 1112 14 34

*11:59-11:604 1 1112 14 41

*11:61 1 1112 w 33

*11:62, 11:684 1 1112 14 35

*11:63, 11:69N4 1 1112 14 42

*11:64-11:654 1 1112 14 36

*11:66, 11:724 1 1112 14 37

*11:70:01-11:70:02 1 1112 14 43

*11:71 1112 14 48

*11:73 1 1112 14 49

*11:74 1 1112 14 33 39

*11:75 1 1112 14 38

*11:76 1 1112 14 44

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



[image: image11.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

HLA-A allele1,2,3

*11:01:01:01-11:01:20, 

11:01:22-11:01:43, 11:01:45-

11:01:56, 11:01:58-11:01:61, 

11:01:63-11:01:66, 11:82-

11:85, 11:87-11:88, 11:91:01-

11:93, 11:95, 11:97, 11:103-

11:104, 11:107, 11:114, 11:116, 

11:119:01-11:119:02, 11:122-

11:123, 11:125, 11:126, 11:132-

11:133, 11:136, 11:138, 11:141, 

11:144-11:146, 11:148-11:152, 

11:160, 11:162, 11:164-11:165, 

11:167, 11:169, 11:177, 11:179, 

11:181, 11:184-11:188, 11:192, 

11:195, 11:197-11:198, 11:200, 

11:202-11:207, 11:212-11:2145

1 1112 14

*11:01:21, 11:01:44, 11:01:62 1 11 14

*11:01:57 1 12 14 41

*11:02:01-11:02:06, 11:77: 

11:113, 11:140, 11:1765

2 1112 14

*11:03 1 3 1112 14 6162

*11:04, 11:130 1 4 11 14

*11:05 1 5 1112 14

*11:06 1 6 1112 14 22

*11:07 1 6 1112 14 16

*11:08 1 7 12 14 23

*11:09, 11:1964 1 8 1112 14

*11:10 1 101112 14 29

*11:11 1 11121314

*11:12 1 1112 15

*11:13, 11:1174 1 1112 14 16

*11:14 2 1112 14 17

*11:15:01 1 1112 14 18

*11:15:02 1 1112 14 28

*11:16 2 1112 14 49

*11:17 1112 2021 27

*11:18 6 1112 22

*11:19, 11:98 1112 21

*11:20 1 1112 14 62

*11:21N 1 1112 14 24

*11:22 1 6 1112 14

*11:23 1 12 14 49

*11:24:01 1 12 14 25

*11:24:02 1 12 1415

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control


[image: image12.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*02:221, 23:41, 24:286, 31:78 8

*02:294, 32:54 5

*02:355 22w

*02:393 23 61

*02:447 6 16 w

*02:453 22 52

*02:484 22 38 42

*02:489 22w 61

*02:527 22 43

*02:541 w 61

*02:557, 24:271 52

*03:01:01:01-03:01:01:03, 

03:01:03-03:01:05, 03:01:07, 

03:01:10-03:01:11, 03:01:13, 

03:01:15-03:01:22, 03:01:24-

03:01:29, 03:01:31-03:01:37, 

03:01:39-03:01:46, 03:01:48, 

03:01:51, 03:01:53-03:01:54, 

03:03N, 03:05:01, 03:06, 

03:11N, 03:13-03:17:02, 03:19-

03:20, 03:25-03:30, 03:33-

03:35, 03:37-03:39, 03:43-

03:49, 03:51-03:53, 03:55-

03:56, 03:58, 03:60-03:62, 

03:67-03:68N, 03:70-03:71, 

03:74, 03:77, 03:79-03:81, 

03:83-03:86, 03:91N-03:94, 

03:96, 03:99-03:104, 03:110, 

03:112, 03:114-03:121, 03:124-

03:126, 03:128-03:132, 03:134, 

03:136-03:139, 03:141, 03:143-

03:148, 03:150-03:151, 03:155-

03:159, 03:162N-03:166, 

03:168N-03:170, 03:175, 

03:181, 03:183-03:186, 03:188-

03:190, 03:192N-03:195, 

03:197N, 03:201-03:202, 

03:207, 03:209, 03:212-03:214, 

03:217

1 4 15 40 62

*03:01:02, 03:05:02 1 4 15 25 40 62

*03:01:06, 03:01:14, 03:01:47, 

03:01:52, 03:54, 03:64, 03:127, 

03:140, 03:200Q

1 4 15 40

*03:01:08, 03:01:12, 03:09, 

03:12, 03:59, 03:72, 03:109, 

03:149, 03:179-03:180, 03:196, 

03:205, 03:216, 03:221

1 4 15 62

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control


[image: image13.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*11:78N 1112 14 38

*11:79, 11:1294 1 1112 14 45

*11:80 1 1112 14 46

*11:81, 11:1084 1 1112 14 52

*11:86, 11:1424 1 1112 14 50

*11:89 1 1112 14 54

*11:90 1 11 14 55

*11:94 1 7 9 12 14 27 56

*11:96, 11:159 1 1112

*11:99N 1 1112 14 56

*11:100 1 1112 14 53

*11:101 2 14

*11:102 1 1112 14 55

*11:105 1 1112 14 38 58

*11:106 1112 44

*11:109N 1 1112 14 49 54

*11:110 2 5 1112 14

*11:111, 11:161, 11:168, 11:190

1112 14

*11:112 1 w12 14 56

*11:115N, 11:180N4 1 1112 14 58

*11:118 1 14 26

*11:120, 11:1714 1 1112 14 63

*11:121 1112 21 43

*11:124, 11:1744 1 1112 14 59

*11:127N, 11:137N4 1 1112 14 60

*11:128 1 1112 14 55 60

*11:131 1 1112 14 20 30

*11:134 1 1112 14 19 42

*11:135 1 1112 14 25 48

*11:139, 11:143 1 14

*11:147 2 1112 14 32

*11:153:01-11:153:02 1 1112 14 54 60

*11:154 1 1112 14 53 63

*11:155 1 7 1112 14 27

*11:156 1 1112 w 30 34

*11:157 1 11w 14

*11:158 1 12 1415 32 37 41 49

*11:163 1 1112 14 45 59

*11:166 1 1112 14 39 54

*11:170Q 1 11 14 18 28

*11:172 1 1112 14 53 55

*11:173, 11:1934 1 1112 14 46 57

*11:175 1 3 1112 14 53 6162

*11:178 1112

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364



[image: image14.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*01:12, 01:19 4 11 21 27 41

*01:13 1 7 12 14 27 40

*01:14, 36:01-36:02, 36:05 4 7 21 27

*01:17 7 12 14 27

*01:21 7 12 21 27 32 41 56

*01:25 7 w12 21 27

*01:26 7 9 12 21 27 56

*01:37 7 12 21 27 53 55

*01:41 21 24

*01:43 7 12 2021 27

*01:45 7 12 21 27 52

*01:51 7 10 12 21 27

*01:56N 7 12 21 27 42

*01:61 3 7 12 21 27

*01:79, 01:104, 01:134 w 7 12 21 27

*01:83:01-01:83:02 7 12 21 27 43

*01:86 7 12 21 27 54 60

*01:101 7 12 21 27 31

*01:109 7 12 21 27 53

*01:115 7 12 21 27 32

*01:125 7 12 21 27 35

*01:126 7 12 21 27 32 41

*01:127 7 1112 21 27

*01:129 7 8 12 21 27

*01:136 9 1112 21 27

*01:143 7 10 12 21 27 31

*01:148 7 12 21 27 55 60

*02:03:01-02:03:08, 02:148, 

02:253, 02:258, 02:264, 02:281, 

02:315, 02:345, 02:370, 02:412, 

02:431, 02:463, 02:466, 02:480, 

02:505, 02:529

w

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control


[image: image15.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*03:01:09, 03:161N, 03:206 4 15 40 62

*03:01:23, 03:08, 03:36N, 

03:57, 03:107-03:108, 03:111, 

03:142, 03:152, 03:172-03:173, 

03:176, 03:178N, 03:199, 

03:203, 03:211, 03:219, 30:89, 

32:04

4 15 62

*03:01:30, 03:01:50, 03:40, 

03:98, 03:105

1 15 40 62

*03:01:38 1 4 15 23 40 62

*03:01:49, 03:75, 03:95, 03:177, 

03:204, 03:215, 74:13

1 15 40

*03:02:01-03:02:04, 03:10, 

03:31, 03:73, 03:76, 03:82, 

03:90, 03:106, 03:160

1 4 11 15 4041

*03:04:01-03:04:03 1 4 15 40 54 60 62

*03:07, 03:133 1 4 15 32 4041

*03:18 1 7 12 15w 27 40

*03:21N 1 4 15 24 40 62

*03:22:01-03:22:02 1 3 4 15 40 62

*03:23:01-03:23:02 1 4 15 40 43 62

*03:24 1 4 15 29 40 62

*03:32 4 11 41

*03:41 4 15 21 62

*03:42 1 3 4 15 32 4041

*03:50 1 4 7 15 40 62

*03:63 1 4 1415 40 62

*03:65, 03:69N 1 4 11 15 40 62

*03:66, 03:153 1 4 7 15 40

*03:78 1 4 15 40 52 62

*03:87 1 4 15 31 40 62

*03:88 1 4 1415 62

*03:89 4 15 43 62

*03:97, 03:122 1 4 15 27 40 62

*03:113 1 4 11 w 4041

*03:123:01 1 4 15 40 49 62

*03:123:02 1 4 15 23 40 49 62

*03:135 1 3 7 12 15w 27 40 56 61

*03:154 1 4 15 32 40 62

*03:167 1 4 11 15 27 40 62

*03:171 1 4 15 32 37 4041 49

*03:174 1 4 15 36 40 62

*03:182 1 4 15 40 53 62

*03:187 1 4 7 15 27 40

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control


[image: image16.png]





1HLA-A*11 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.20.0, April 2015.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 5: Specific PCR product of 180 bp in the A*11:05 allele. Specific PCR product of 210 bp in the A*11:110 and the A*02:294, 32:54 and 66:08 alleles.
Primer mix 6: Specific PCR product of 190 bp in the A*11:07 and 11:22 and the A*01:79w alleles. Specific PCR product of 550 bp in the A*11:06 and 11:18 and the A*01:104w, 01:134w, 02:38, 02:101:01, 02:135, 02:154, 02:447, 02:543, 25:11, 26:03:01, 26:06, 26:21, 26:30, 26:36, 26:92 and 26:111 alleles.

Primer mix 8: Specific PCR fragment of 100 bp in the A*11:09 and 11:34 and the A*26:75 alleles. Specific PCR fragment of 175 bp in the A*11:196 and 11:201 and the A*01:129, 02:221, 23:41, 24:286, 29:15 and 31:78 alleles.  
Primer mix 10: Specific PCR fragment of 120 bp in the A*11:10 and the A*01:51, 25:01:01-25:12N, 25:14-25:31, 26:01:01-26:06, 26:08-26:18, 26:20-26:29, 26:32-26:34, 26:36-26:43:02, 26:45-26:64, 26:66-26:71N, 26:73-26:75, 26:77-26:91, 26:93-26:111, 34:01:01-34:08, 34:10N-34:12, 66:01:01-66:01:03, 66:04-66:11, 66:13-66:15, 66:17-66:20, 66:22 and 69:02 and in the B*07:102, B*08:18, B*35:178, 35:282, B*73:01-73:02 and C*06:147 alleles. Specific PCR fragment of 280 bp in the A*11:43 and the A*01:143, 29:66 and 31:03 alleles. Specific PCR fragment of 120 and 280 bp in the A*26:35 and 33:13 alleles.
Primer mix 15: Specific PCR fragment of 100 bp in the A*11:12 and the A*02:34w-02:35:03w, 02:56:01-02:56:02, 02:62, 02:78, 02:103, 03:01:02, 03:01:06, 03:02:01-03:02:04, 03:05:02-03:06, 03:10, 03:18, 03:31, 03:50, 03:65-03:66, 03:69N, 03:73, 03:75-03:76, 03:82, 03:90, 03:95, 03:106, 03:113w, 03:135, 03:150, 03:153, 03:160, 03:167, 03:177, 03:187, 03:200, 03:204, 03:208, 03:215, 03:218, 24:24, 24:290, 29:01:01:01-29:03, 29:04w, 29:05-29:18, 29:21-29:29, 29:31-29:33, 29:35-29:47, 29:49-29:65, 29:67-29:73, 29:75-29:78N, 30:17, 30:52, 30:72, 31:89, 32:17, 66:02, 66:12, 66:16, 66:21, 68:01:01:01-68:01:17, 68:01:19-68:02:09, 68:06-68:14, 68:16-68:19, 68:21:01-68:23, 68:25-68:30, 68:32-68:39, 68:41-68:71, 68:73-68:84, 68:86-68:89, 68:91-68:92, 68:94N-68:96, 68:98-68:108, 68:110-68:122, 68:124-68:126, 68:127w, 68:128-68:132, 69:01, 69:03 and 74:13  alleles. Specific PCR fragment of 215 bp in the A*11:24:02-11:25, 11:31, 11:35 and 11:158 and the A*01:06, 02:156, 02:338, 03:01:10, 03:08, 03:29-03:30, 03:34, 03:36, 03:41, 03:62-03:63, 03:88-03:89, 03:107-03:108, 03:152, 03:172, 03:176, 30:02:01-30:02:04, 30:02:06-30:03, 30:07 30:09-30:10, 30:12-30:13, 30:22, 30:27N, 30:31-30:34, 30:44-30:45, 30:50-30:51, 30:57, 30:61, 30:64, 30:66-30:70, 30:76N, 30:84-30:85, 31:03-31:04, 32:04, 33:49, 34:02:01, 34:02:03-34:04, 34:07-34:10N and 74:23 alleles. Specific PCR fragment of 100 and 215 bp in the A*03:01:01:01-03:01:01:03, 03:01:03-03:01:05, 03:01:07-03:01:09, 03:01:11-03:01:48, 03:01:50-03:01:54, 03:03N-03:05:01, 03:07, 03:09, 03:11N-03:17:02, 03:19-03:28, 03:33, 03:35, 03:37-03:40, 03:42-03:49, 03:51-03:61, 03:64, 03:67-03:68N, 03:70-03:72, 03:74, 03:77-03:81, 03:83-03:87, 03:91N-03:94, 03:96-03:105, 03:109-03:112, 03:114-03:134, 03:136-03:149, 03:151, 03:154-03:159, 03:161N-03:166, 03:168N-03:171, 03:173-03:175, 03:178N-03:186, 03:188-03:197N, 03:199, 03:201-03:203, 03:205-03:207, 03:209-03:214, 03:216-03:217, 03:219, 03:221, 30:01:01-30:01:10, 30:08, 30:11:01-30:11:02, 30:14L-30:16, 30:18-30:20, 30:23-30:26, 30:30, 30:35-30:43, 30:48-30:49, 30:53-30:56, 30:58-30:60, 30:62-30:63, 30:65, 30:71, 30:73N-30:75, 30:78N-30:79, 30:81-30:83, 30:86-30:89, 30:91-30:94 and 34:09  alleles. Both PCR fragments may not always be obtained.
Primer mix 16: Specific PCR fragment of 95 bp in the A*11:13 allele. Specific PCR fragment of 210 bp in the A*11:07, 11:26, 11:54 and 11:117 and the A*02:38, 02:101:01, 02:154, 02:447, 02:543, 03:18w, 03:135w, 23:10w, 24:10:01-24:10:02, 24:46w, 24:210, 24:300, 25:10, 31:24 and 33:61 alleles.

Primer mix 18: Specific PCR product of 225 bp in the A*11:47 allele. Specific PCR product of 240 bp in the A*11:15:01 and 11:170Q and the A*30:39 alleles. 
Primer mix 19: Specific PCR product of 100 bp in the A*11:134 allele.  Specific PCR product of 175 bp in the A*11:42 and in the B*39:75 and B*40:268 alleles. Specific PCR product of 265 bp in the A*11:46 and 11:215N and in the B*07:111N alleles.

Primer mix 20: Specific PCR fragment of 180 bp in the A*11:42 and 11:131 and in the B*39:75 and B*40:268 alleles. Specific PCR fragment of 235 bp in the A*11:17 and 11:40 and the A*01:43, 26:01:01-26:01:29, 26:01:31-26:02:02, 26:04, 26:07:01-26:15, 26:17-26:19, 26:22-26:29, 26:31-26:34, 26:36-26:43:02, 26:45-26:56, 26:58-26:77, 26:79-26:91, 26:93-26:110 and 43:01 alleles.

Primer mix 22: Specific PCR fragment of 150 bp in the A*11:37 and the B*27:87 alleles. Specific PCR fragment of 230 bp in the A*11:06 and 11:18 and the A*02:05:01-02:06:04, 02:06:06-02:06:18, 02:08, 02:10, 02:14, 02:21, 02:28, 02:41, 02:44, 02:51, 02:54, 02:57, 02:61, 02:72, 02:79:01-02:79:02, 02:84, 02:91, 02:99, 02:106, 02:108, 02:122, 02:126-02:127, 02:137, 02:142-02:144, 02:154, 02:169-02:170, 02:172, 02:178-02:180, 02:229, 02:232, 02:244, 02:248, 02:259, 02:271, 02:278, 02:286, 02:290, 02:295, 02:300, 02:310, 02:324, 02:328, 02:330, 02:333, 02:337, 02:344, 02:355, 02:358-02:359, 02:373N, 02:376, 02:382, 02:387, 02:398, 02:404-02:405, 02:409, 02:413, 02:415, 02:419-02:421, 02:428, 02:433, 02:438, 02:453-02:454, 02:465, 02:470-02:476N, 02:484, 02:489, 02:493, 02:495-02:496, 02:506N-02:507, 02:527, 02:532, 02:546, 02:550, 02:558, 02:572, 26:03:01, 26:06, 26:21, 26:30, 26:78, 26:92, 26:111 and 68:15 alleles.

Primer mix 24: Specific PCR fragment of 295 bp in the A*11:48 and the A*01:41 alleles. Specific PCR fragment of 465 bp in the A*11:21N and the A*01:04N, 03:21N, 23:07N and 24:11N alleles.  
Primer mix 25: Specific PCR fragment of 190 bp in the A*11:135 and 03:210 alleles. Specific PCR fragment of 220 bp in the A*11:24:01 and 11:47 and the A*03:01:02, 03:05:02, 30:02:05, 34:02:02 and 80:01:01:01-80:03 alleles. 

Primer mix 26: Specific PCR fragment of 305 bp in the A*11:26 and 11:118 and the A*01:03, 26:32, 26:70, 29:66, 32:62, 33:13, 36:03, 74:10 and in the C*06:72 alleles. Specific PCR fragment of 370 bp in the A*11:41 and the A*34:12 and in the C*06:90 and C*07:277alleles.

Primer mix 27: Specific PCR fragment of 160 bp in the A*11:208N and the A*26:71N and 68:94N alleles. Specific PCR fragment of 205 bp in the A*11:17, 11:155 and the A*01:01:66, 01:72, 01:102, 01:167, 03:187, 36:01-36:02 and 36:04-36:05 alleles. Specific PCR fragment of 260 bp in the A*11:27, 11:38-11:39, 11:55, 11:94, 11:209, and the A*01:01:12, 01:01:16, 01:01:55, 01:06, 01:10, 01:12, 01:19, 01:21, 01:26, 01:119, 01:126, 01:136, 01:160N, 03:97, 03:122, 03:135, 03:167, 24:92, 30:75 and 80:01:01:01-80:03 alleles. Specific PCR fragment of 205 and 260 bp in the A*01:01:01:01-01:01:11, 01:01:13-01:01:15, 01:01:17-01:01:54, 01:01:56-01:01:57, 01:01:59-01:01:65, 01:01:67-01:02, 01:04N, 01:07-01:09, 01:11N, 01:13-01:18N, 01:22N-01:25, 01:27N-01:33, 01:35-01:40, 01:42-01:62, 01:64-01:65, 01:67:01-01:71, 01:74-01:88, 01:90-01:118, 01:120-01:125, 01:127-01:129, 01:131-01:135, 01:137-01:159, 01:161-01:166, 01:168-01:175, 01:178N-01:179N and 03:18 alleles.
Primer mix 28: Specific PCR fragment of 240 bp in the A*11:15:02, 11:49 and 11:170Q alleles. Specific PCR fragment of 280 bp in the A*26:35 allele.

Primer mix 30: Specific PCR fragment of 180 bp in the A*11:131 allele. Specific PCR fragment of 200 bp in the A*11:156 allele. Specific PCR fragment of 235 bp in the A*11:45, 11:50Q and 11:182Q and the A*66:11 alleles. Specific PCR fragment of 270 bp in the A*11:29, 11:55, 11:215N and in the B*07:111N alleles.

Primer mix 31: Specific PCR fragment of 230 bp in the A*11:49 allele. Specific PCR fragment of 285 bp in the A*11:30 and 11:43 and the A*01:101, 01:143, 03:87, 29:66, 30:92, 31:03 and 33:13 alleles.

Primer mix 32: Specific PCR fragment of 205 bp in the A*11:31, 11:35, 11:158 and 11:183 and the A*01:21, 01:126, 02:156, 02:338, 03:07, 03:42, 03:133, 03:171, 30:04:01-30:04:02, 30:06, 30:09, 30:17, 30:29, 30:46, 30:77, 30:90, 31:03-31:04, 33:49, 68:103:01-68:103:02 and 74:23 alleles. Specific PCR fragment of 295 bp in the A*11:32, 11:50Q and 11:147 and the A*01:115, 03:154 and 30:51 alleles.

Primer mix 33: Specific PCR fragment of 110 bp in the A*11:61 allele. Specific PCR fragment of 280 bp in the A*11:74 allele. Specific PCR fragment of 305 bp in the A*11:56 and in the C*07:221 alleles.

Primer mix 34: Specific PCR fragment of 85 bp in the A*11:58 allele. Specific PCR fragment of 170 bp in the A*11:67 and 11:208N and the A*26:71N and 68:94N alleles. Specific PCR fragment of 200 bp in the A*11:156 allele. Specific PCR fragment of 265 bp in the A*11:46 allele.

Primer mix 35: Specific PCR fragment of 105 bp in the A*11:62 and the A*01:125 allele. Specific PCR fragment of 290 bp in the A*11:68 allele. 

Primer mix 36: Specific PCR fragment of 65 bp in the A*11:65 and in the A*03:174 and in the B*44:36 alleles. Specific PCR fragment of 230 bp in the A*11:210N allele. Specific PCR fragment of 380 bp in the A*11:64 and the A*31:65 alleles. 

Primer mix 37: Specific PCR fragment of 105 bp in the A*11:66 allele. Specific PCR fragment of 185 bp in the A*11:57 and 11:158 and the A*02:156, 02:338, 03:171 and 68:103:01-68:103:02 alleles. Specific PCR fragment of 290 bp in the A*11:72 allele. 

Primer mix 38: Specific PCR fragment of 95 bp in the A*11:105 and 11:189 and the A*02:484, 26:77 and 66:13 alleles. Specific PCR fragment of 120 bp in the A*11:75 and in the B*07:117 alleles. Specific PCR fragment of 225 bp in the A*11:78N allele. 

Primer mix 39: Specific PCR fragment of 205 bp in the A*11:166 and 80:01:01:01-80:03 and in the C*04:129, C*05:25, 05:42, C*06:05, 06:67, C*07:101, 07:148, 07:161 and C*08:28 alleles. Specific PCR fragment of 280 bp in the A*11:52Q and 11:74 alleles. 
Primer mix 41: Specific PCR fragment of 90 bp in the A*11:59 and the A*03:191 alleles. Specific PCR fragment of 215 bp in the A*11:01:57, 11:31, 11:35, 11:60, 11:158, 11:183 and 11:209 and the A*01:12, 01:19, 01:21, 01:126, 02:156, 02:338, 03:02:01-03:02:04, 03:07, 03:10, 03:31-03:32, 03:42, 03:73, 03:76, 03:82, 03:90, 03:106, 03:113, 03:133, 03:160, 03:171, 03:198, 03:218, 24:92, 30:04:01w-30:04:02w, 30:06w, 30:09, 30:17w, 30:29w, 30:46w, 30:77w, 30:90w, 31:03-31:04, 33:49, 68:103:01w-68:103:02w and 74:23 alleles. 

Primer mix 42: Specific PCR fragment of 95 bp in the A*11:134 and 11:189 and the A*02:484 and 26:77 alleles. Specific PCR fragment of 175 bp in the A*11:63 allele. Specific PCR fragment of 560 bp in the A*11:69N and the A*01:56N alleles.

Primer mix 44: Specific PCR fragment of 145 bp in the A*11:76 allele. Specific PCR fragment of 215 bp in the A*11:106 allele.

Primer mix 45: Specific PCR fragment of 210 bp in the A*11:79 allele. Specific PCR fragment of 340 bp in the A*11:129 and 11:163 alleles. Specific PCR fragment of 400 bp in the A*11:210N allele.   
Primer mix 46: Specific PCR fragment of 75 bp in the A*11:193 allele. Specific PCR fragment of 120 bp in the A*11:80 and in the C*16:79 allele. Specific PCR fragment of 175 bp in the A*11:173 allele. 
Primer mix 49: Specific PCR fragment of 180 bp in the A*11:16, 11:35, 11:57, 11:73, 11:109N and 11:158 and the A*02:156, 02:338, 03:123:01-03:123:02, 03:171 and 68:103:01-68:103:02 alleles. Specific PCR fragment of 230 bp in the A*11:23, 11:50Q and 11:182Q alleles.

Primer mix 50: Specific PCR fragment of 90 bp in the A*11:34 and the A*26:75 alleles. Specific PCR fragment of 210 bp in the A*11:142 allele. Specific PCR fragment of 310 bp in the A*11:86 and A*03:218 alleles. Specific PCR fragment of 400 bp in the A*11:210N alleles. 
Primer mix 52: Specific PCR fragment of 230 bp in the A*11:81 allele. Specific PCR fragment of 545 bp in the A*11:108 and the A*01:45, 02:453, 02:557, 03:78, 24:271 and 66:17 alleles.

Primer mix 53: Specific PCR fragment of 150 bp in the A*11:154 and 11:172 and the A*01:37 alleles. Specific PCR fragment of 195 bp in the A*11:100 and 11:175 and the A*01:109 and 03:182 alleles.  
Primer mix 54: Specific PCR fragment of 200 bp in the A*11:109N and 11:166 and the A*80:01:01:01-80:03 and in the C*04:129, C*05:25, 05:42, C*06:05, 06:67, C*07:101, 07:148, 07:161 and C*08:28 alleles. Specific PCR fragment of 270 bp in the A*11:89, 11:153:01-11:153:02 and the 01:86, 03:04:01-03:04:03 and 30:34 alleles.

Primer mix 55: Specific PCR fragment of 115 bp in the A*11:102 and 11:128 and the A*01:148 and 24:249 alleles. Specific PCR fragment of 145 bp in the A*11:172 and the A*01:37 alleles. Specific PCR fragment of 170 bp in the A*11:194 allele.  Specific PCR fragment of 235 bp in the A*11:90 and the A*03:208 and 80:01:01:01-80:02 alleles.

Primer mix 56: Specific PCR fragment of 65 bp in the A*11:94, 11:112 and 11:211 and the A*01:10, 01:21, 01:26 and 03:135 and in the B*15:90w and B*45:05w alleles. 

Specific PCR fragment of 230 bp in the A*11:99N allele. 

Primer mix 57: Specific PCR fragment of 75 bp in the A*11:193 alleles.  Specific PCR fragment of 175 bp in the A*11:173 allele. Specific PCR fragment of 245 bp in the A*11:36 allele. 

Primer mix 58: Specific PCR fragment of 95 bp in the A*11:105 and 11:115N and the A*66:13 alleles. Specific PCR fragment of 245 bp in the A*11:180N allele. 
Primer mix 59: Specific PCR fragment of 165 bp in the A*11:174 allele. Specific PCR fragment of 340 bp in the A*11:163 allele. Specific PCR fragment of 370 bp in the A*11:124 allele. 

Primer mix 60: Specific PCR fragment of 105 bp in the A*11:128, 11:137N and the A*01:148 alleles. Specific PCR fragment of 270 bp in the A*11:127N and 11:153:01- 11:153:02 and the A*01:86, 03:04:01-03:04:03 and 30:34 alleles.
Primer mix 61: Specific PCR fragment of 130 bp in the A*11:03 and 11:175 and the A*02:393, 02:489, 02:541, 03:135, 24:177, 24:289, 26:28 and 26:52 and in the B*15:173, B*50:11 and C*12:74 alleles. Specific PCR fragment of 265 bp in the A*11:54 and the A*33:70 alleles. 

Primer mix 63: Specific PCR fragment of 145 bp in the A*11:171 alleles. Specific PCR fragment of 170 bp in the A*11:154 and 11:194 alleles. Specific PCR fragment of 275 bp in the A*11:120 allele.  
4The following HLA-A*11 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

	Alleles
	Primer mix
	Alleles
	Primer mix

	A*11:09, 11:196
	8
	A*11:79, 11:129
	45

	A*11:13, 11:54
	16
	A*11:81, 11:108
	52

	A*11:58, 11:67
	34
	A*11:86, 11:142
	50

	A*11:59, 11:60
	41
	A*11:115N, 11:180N
	58

	A*11:62, 11:68
	35
	A*11:120, 11:171
	63

	A*11:63, 11:69N
	42
	A*11:124, 11:74
	59

	A*11:64, 11:65 
	36
	A*11:127N, 11:137N
	60

	A*11:66, 11:72
	37
	A*11:173, 11:193
	46, 57


The HLA-A*11 subtyping kit cannot distinguish the silent mutations in the A*11:01:01-11:01:20, 11:01:22-11:01:43, 11:01:45-11:01:56, 11:01:58-11:01:61 and 11:01:63-11:01:66 alleles, the A*11:01:21, 11:01:44, 11:01:62 alleles, the A*11:02:01-11:02:06 alleles, the A*11:33:01-11:33:02 alleles, the A*11:70:01-11:70:02 alleles or the A*11:153:01-11:153:02 alleles. 

5This lot of the HLA-A*11 subtyping kit cannot distinguish the A*11:01:01:01-11:01:20, 11:01:22-11:01:43, 11:01:45-11:01:56, 11:01:58-11:01:61, 11:01:63-11:01:66, 11:82-11:85, 11:87-11:88, 11:91:01-11:93, 11:95, 11:97, 11:103-11:104, 11:107, 11:114, 11:116, 11:119:01-11:119:02, 11:122-11:123, 11:125, 11:132-11:133, 11:136, 11:138, 11:141, 11:144-11:146, 11:148-11:152, 11:160, 11:162, 11:164-11:165, 11:167, 11:169, 11:177, 11:179, 11:181, 11:184-11:188, 11:192, 11:195, 11:197-11:198, 11:200, 11:202-11:207, 11:212-11:214 alleles and the A*11:126 allele, nor the A*11:02:01-11:02:06, 11:113, 11:140, 11:176 alleles and the A*11:77 allele.
‘w’, may be weakly amplified.
Change in revision R01 compared to R00:

1. Primer mix 5 does not amplify the A*11:77 and 11:126 and the A*03:26 alleles. This has been corrected in the Specificity and Interpretation Tables. Thus, this lot of the HLA-A*11 subtyping kit cannot distinguish the A*11:01:01:01-11:01:20, 11:01:22-11:01:43, 11:01:45-11:01:56, 11:01:58-11:01:61, 11:01:63-11:01:66, 11:82-11:85, 11:87-11:88, 11:91:01-11:93, 11:95, 11:97, 11:103-11:104, 11:107, 11:114, 11:116, 11:119:01-11:119:02, 11:122-11:123, 11:125, 11:132-11:133, 11:136, 11:138, 11:141, 11:144-11:146, 11:148-11:152, 11:160, 11:162, 11:164-11:165, 11:167, 11:169, 11:177, 11:179, 11:181, 11:184-11:188, 11:192, 11:195, 11:197-11:198, 11:200, 11:202-11:207, 11:212-11:214 alleles and the A*11:126 allele, nor the A*11:02:01-11:02:06, 11:113, 11:140, 11:176 alleles and the A*11:77 allele.

Change in revision R02 compared to R01:

1. Primer mix 46 does not amplify the A*11:100 and 11:175 and the A*01:109 and 03:182 alleles. This has been corrected in the Specificity and Interpretation Tables. 

Olerup SSP AB

For In Vitro Diagnostic Use 
Franzengatan 5


SE-112 51 Stockholm

Sweden


Page 1 of 18

