Gel Documentation Form and Worksheet

HLA-A*26 (101.424-12/12u) 	Lot No: 6G1	Expiry Date: 2022-09-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________	Failed lanes: ___________ 	 Comments:_________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 1 and 39 have a tendency to giving rise to primer oligomer formation.

Primer mixes 4, 12, 20, 22, 26 and 33 may have tendencies of unspecific amplifications.

Primer mixes 7 and 25 may give rise to a lower yield of HLA-specific PCR product than the other A*26 primer mixes.

In primer mix 4 the positive control band may be weaker than for other HLA-A*26 primer mixes.

Primer mix 48 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pair.

1HLA-A*26 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.26.0, October 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-A*26 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-A*26 alleles | Other amplified HLA-A alleles |
| --- | --- | --- | --- |
| 9 | 145 bp 190 bp | *26:12, 26:18 *26:29, 26:49 | *02:309, 02:454, 25:28, 25:30, 31:03-31:04, 31:123, 34:06, 66:06, 74:01:03 *01:01:56, 25:36, 66:10 |
| 10 | 85 bp 260 bp | *26:07:01-26:07:02 *26:20 | *01:83:01-01:83:02, 02:146 *01:01:13, 02:146w, 02:644w |
| 12 | 140 bp 240 bp | *26:14, 26:18, 26:28, 26:52, 26:73, 26:112, 26:146 *26:16 | *03:01:19, 03:103:02, 25:09, 31:03-31:04, 31:123, 34:03, 34:06, 34:17, 66:22, 74:01:03 *01:02, 01:20, 01:188, 01:190, 24:04, 24:109, 24:129, 29:37, 29:56, 30:57, 30:90 |
| 13 | 110 bp 255 bp | *26:45, 26:106 *26:17 | *01:184 |
| 17 | 105 bp 175 bp | *26:74 *26:09, 26:91 | *02:309, 02:454, 03:01:19, 03:103:02, 25:06, 31:03-31:04, 31:123, 34:01:01:01-34:09, 34:12-34:17, 74:01:03 |
| 18 | 125 bp 205 bp | *26:31 *26:21, 26:115 | *01:60, C*06:71, C*07:581 |
| 19 | 190 bp 245 bp | *26:40 *26:22 | *01:20w, 01:66w, 01:130w, 02:38, 02:101:01, 02:447, 02:543, 66:09 |
| 20 | 60 bp 210 bp | *26:27 *26:23 | *01:243w *01:182w, 11:185 |
| 21 | 115 bp 135 bp 205 bp | *26:24, 26:99 *26:82 *26:41 | *02:140, 02:241, 02:684, 31:48, 33:15, 33:31 |
| 22 | 100 bp 170 bp 220 bp | *26:25N *26:98 *26:38 | C*04:199, C*06:75 |
| 23 | 130 bp 165 bp | *26:56, 26:82 *26:26 |  |
| 27 | 125 bp 150 bp 275 bp | *26:76 *26:145N *26:35 |  |
| 28 | 90 bp 265 bp | *26:03:01, 26:06, 26:21, 26:36, 26:78, 26:92, 26:111, 26:146 *26:116 | *01:104, 01:134, 01:229, 02:346, 02:427, 11:06, 25:11, 80:01:01:01w |
| 29 | 330 bp 475 bp | *26:37 *26:104 | *24:74:02 |
| 30 | 120 bp 235 bp | *26:15, 26:29, 26:56, 26:99 *26:33 | *02:140, 11:91:01, 31:48, 33:15, B*35:108:01, B*44:248, B*53:26, C*12:91 |
| 32 | 90 bp 155 bp 205 bp | *26:61 *26:43:01 *26:21, 26:115 | *02:309, 02:454, 03:01:19, 03:103:02, 25:19:01-25:19:02, 25:30, 31:03-31:04, 31:123, 34:02:01, 34:02:03-34:04, 34:06-34:09, 34:13, 34:15, 66:06, 74:01:03 |
| 33 | 105 bp 170 bp 200 bp | *26:46 *26:98 *26:53 |  |
| 34 | 140 bp 190 bp | *26:47, 26:127N *26:29, 26:49 | *25:08, 66:18 *01:01:56, 25:36, 66:10 |
| 35 | 115 bp 245 bp | *26:48, 26:69, 26:114 *26:59 | *34:01:01:01-34:01:02, 34:05, 34:11-34:12, 34:14, 34:16-34:17, C*04:49, C*07:262, C*07:417, C*07:475 |
| 36 | 125 bp 150 bp 215 bp | *26:54 *26:127N *26:55 | *02:454, 25:05 |
| 37 | 245 bp 415 bp | *26:50 *26:64, 26:66 | *68:89 |
| 38 | 90 bp 150 bp 190 bp | *26:75 *26:145N *26:51 | *01:205, 11:34 *01:168, 11:262 |
| 39 | 220 bp 410 bp | *26:63 *26:62, 26:64, 26:72 | *24:181 *02:609, 23:09, 24:129, 68:89 |
| 40 | 90 bp 150 bp 255 bp | *26:57 *26:42 *26:57 | *68:86, B*39:104 *68:86, B*39:104 |
| 42 | 425 bp 460 bp | *26:66 *26:58 | *02:81, 02:87, 02:112, 02:124, 02:129, 02:136, 02:571, 23:01:01:01-23:68, 23:70-23:88, 24:02:01:01-24:02:32, 24:02:34-24:03:04, 24:05:01-24:11N, 24:13:01-24:15, 24:17-24:18, 24:20:01:01-24:25, 24:27, 24:29-24:43, 24:45N-24:64, 24:66-24:88, 24:90:01N-24:99, 24:101-24:108, 24:110-24:128, 24:130-24:183N, 24:185N-24:210, 24:212-24:213, 24:215-24:239, 24:242-24:289, 24:291-24:293, 24:295-24:372, 24:374-24:400, 25:01:01:01-25:46, 68:36, B*27:163 |
| 43 | 90 bp 175 bp | *26:107N *26:103 | *25:18 |
| 44 | 105 bp 235 bp | *26:68 *26:130 | *25:27:01, 66:14 |
| 45 | 155 bp 275 bp | *26:42, 26:71N *26:104 | *11:208N, 68:94N, B*35:173:01N |

4The following HLA-A*26 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| A*26:17, 26:106 | 13 | A*26:46, 26:53 | 33 |
| A*26:23, 26:27 | 20 | A*26:54-26:55 | 36 |
| A*26:24, 26:41 | 21 | A*26:59, 26:69 | 35 |
| A*26:25N, 26:38 | 22 | A*26:62-26:63 | 39 |
| A*26:36, 26:116 | 28 | A*26:68, 26:130 | 44 |
| A*26:43:01, 26:61 | 32 | A*26:103, 26:107N | 43 |

‘w’, may be weakly amplified

Changes in revision R01 compared to R00:

1.	The expiration date has been altered due to extension of shelf-life.