Gel Documentation Form and Worksheet

HLA-A*29 (101.428-12/12u) 	Lot No: 9F9	Expiry Date: 2020-06-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 9, 12, 13, 16 and 20 may have tendencies of unspecific amplifications.

Primer mixes 10 and 19 have a tendency giving rise to primer oligomer formation.

Primer mix 13 may give rise to a lower yield of HLA-specific PCR product than the other A*29 primer mixes.

Primer mix 28 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-A*29 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.29.0, August 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-A*29 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-A*29 alleles | Other amplified HLA-A alleles |
| --- | --- | --- | --- |
| 4 | 110 bp 165 bp | *29:35, 29:51, 29:69, 29:73 *29:03, 29:33 | *02:24:02, 02:507 *23:03:01, 24:21:03, 24:208, 31:05, 32:13, 33:10 |
| 5 | 130 bp 185 bp | *29:04 *29:23 |  |
| 6 | 105 bp 130 bp | *29:19 *29:05, 29:33, 29:40, 29:77, 29:87 | *11:01:28, 11:01:77, 24:21:03, 24:208, 31:24, 32:02, 32:22, 33:59, 33:102 |
| 7 | 105 bp 210 bp 475 bp | *29:19-29:20, 29:34 *29:06 *29:82 | *31:51, 32:12, 74:26, B*08:01:07w, B*15:02:07w, B*15:17:03, B*27:07:05, B*41:01:05w, B*55:02:10, C*02:02:15w, C*03:03:20w, C*04:175w, C*08:01:15w |
| 8 | 85 bp 160 bp | *29:07, 29:49 *29:46 | *11:139, 23:53, 23:70, 24:17, 24:41, 24:208 |
| 9 | 80 bp 170 bp | *29:16 *29:08N | *01:157, 03:27, 11:233 |
| 10 | 90 bp 165 bp 215 bp | *29:09, 29:33, 29:51, 29:73 *29:17, 29:43 *29:54 | *02:24:02, 02:507, 03:01:18, 11:01:28, 11:01:77, 24:21:03, 24:208, 31:24, 32:33:01, 33:34 |
| 11 | 110 bp 190 bp | *29:14, 29:35 *29:10:01-29:10:02, 29:23 | C*08:01:15 |
| 12 | 80 bp 145 bp | *29:02:04, 29:18, 29:48 *29:78N | *03:01:39, 32:01:01:01-32:01:07, 32:01:09-32:01:17, 32:01:19-32:03, 32:05-32:55:02, 32:57-32:69, 32:71, 32:73-32:103, 68:01:28, 74:01:01-74:13, 74:15-74:28 |
| 13 | 90 bp 165 bp 260 bp | *29:11, 29:51, 29:73 *29:12, 29:92 *29:55 | *02:24:02, 02:507 *31:16, 33:58 |
| 16 | 95 bp 160 bp 190 bp | *29:15 *29:21, 29:43 *29:53 | *02:221, 23:41, 31:78 |
| 17 | 100 bp 130 bp 190 bp 215 bp | *29:51, 29:69, 29:73 *29:24, 29:40 *29:27, 29:53 *29:54 | *02:24:02, 02:507 |
| 18 | 225 bp 260 bp | *29:37, 29:56 *29:36 | *32:07, 33:119 |
| 19 | 160 bp 260 bp 505 bp | *29:25 *29:55 *29:26 |  |
| 22 | 115 bp 260 bp | *29:29 *29:13 | *01:148, 11:128, 26:85, 68:58:01-68:58:02 *24:82, 31:07-31:08, 31:10, 32:42 |
| 23 | 75 bp 510 bp | *29:32 *29:81 | *02:24:02, 02:65, 02:152, 02:507, 23:03:01, 23:83, 24:21:03, 31:01:02:01-31:02, 31:05, 31:07-31:61, 31:63-31:66, 31:70-31:119, 31:121-31:127, 32:01:01:01-32:01:06, 32:01:08-32:01:11, 32:01:13-32:03, 32:05-32:27N, 32:29, 32:31, 32:33:01, 32:34-32:47, 32:49-32:65, 32:67-32:93, 32:95-32:100, 32:102-32:103, 33:01:01:01-33:01:04, 33:01:06-33:01:10, 33:03:01-33:03:18, 33:03:20-33:03:23, 33:03:25-33:03:26, 33:03:28-33:17, 33:20-33:31, 33:33-33:37, 33:39-33:130, 74:01:01-74:28, B*15:17:03 |

4The following HLA-A*29 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| A*29:06, 29:34 | 7 |
| A*29:07, 29:46 | 8 |
| A*29:11, 29:92 | 13 |
| A*29:20, 29:82 | 7 |

‘w’, might be weakly amplified.