# Gel Documentation Form and Worksheet

# HLA-A*29 (101.428-12/12u) 	Lot No: 3H6	Expiry Date: 2023-03-01

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

# Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

## PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 9, 12, 13, 16 and 20 may have tendencies of unspecific amplifications.

Primer mixes 10 and 19 have a tendency giving rise to primer oligomer formation.

Primer mix 13 may give rise to a lower yield of HLA-specific PCR product than the other A*29 primer mixes.

Primer mix 30 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

Change in revision R01 compared to R00:

- The expiration date has been altered due to extension of shelf-life.

1HLA-A*29 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.29.0, August 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-A*29 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-A*29 alleles | Other amplified HLA-A alleles |
| --- | --- | --- | --- |
| 4 | 110 bp 165 bp | *29:35, 29:51, 29:69, 29:73 *29:03, 29:33 | *02:24:02, 02:507, 32:109 *23:03:01, 24:21:03, 24:208:01, 31:05, 32:13, 33:10 |
| 5 | 130 bp 185 bp | *29:04 *29:23 |  |
| 8 | 85 bp 160 bp | *29:07, 29:49 *29:46 | *11:139, 23:53, 23:70, 24:17, 24:41, 24:208:01-24:208:02 |
| 9 | 80 bp 170 bp | *29:16 *29:08N | *01:157, 03:27, 11:233 |
| 10 | 90 bp 165 bp 215 bp | *29:09, 29:33, 29:51, 29:73 *29:17, 29:43 *29:54 | *02:24:02, 02:507, 03:01:18, 11:01:28, 11:01:77, 24:21:03, 24:208:01, 31:24, 32:33:01, 32:109, 33:34 |
| 11 | 110 bp 190 bp | *29:14, 29:35 *29:10:01-29:10:02, 29:23 | C*07:04:13, C*08:01:15 |
| 12 | 80 bp 145 bp | *29:01:11, 29:02:04, 29:18, 29:48 *29:78N | *03:01:39, 32:01:01:01-32:01:07, 32:01:09-32:01:17, 32:01:19-32:01:29, 32:01:31-32:03, 32:05-32:55:02, 32:57-32:69, 32:71, 32:73-32:107, 32:109-32:112N, 32:114-32:118, 68:01:28, 68:01:36, 74:01:01-74:01:05, 74:01:07-74:13, 74:15-74:32N |
| 13 | 90 bp 165 bp 260 bp | *29:11, 29:51, 29:73 *29:12, 29:92 *29:55 | *02:24:02, 02:507, 32:109 *31:16, 33:58 |
| 16 | 95 bp 160 bp 190 bp | *29:15 *29:21, 29:43 *29:53 | *02:221, 23:41, 31:78 |
| 17 | 100 bp 130 bp 190 bp 215 bp | *29:51, 29:69, 29:73 *29:24, 29:40 *29:27, 29:53 *29:54 | *02:24:02, 02:507, 32:109 |
| 18 | 225 bp 260 bp | *29:37, 29:56 *29:36 | *32:07, 33:119 |
| 19 | 160 bp 260 bp 505 bp | *29:25 *29:55 *29:26 |  |
| 22 | 115 bp 260 bp | *29:29 *29:13 | *01:148, 11:128, 26:85, 33:139, 68:58:01-68:58:02 *24:82, 31:07-31:08, 31:10, 32:42 |
| 23 | 75 bp 510 bp | *29:32 *29:81 | *02:24:02, 02:65, 02:152, 02:507, 23:03:01, 23:83, 24:21:03, 31:01:02:01-31:02, 31:05, 31:07-31:61, 31:63-31:66, 31:70-31:119, 31:121-31:146, 32:01:01:01-32:01:06, 32:01:08-32:01:11, 32:01:13-32:01:27, 32:01:29-32:03, 32:05-32:27N, 32:29, 32:31, 32:33:01, 32:34-32:47, 32:49-32:65, 32:67-32:93, 32:95-32:100, 32:102-32:118, 33:01:01:01-33:01:04, 33:01:06-33:01:12, 33:03:01:01-33:03:18, 33:03:20-33:03:23, 33:03:25-33:03:26, 33:03:28-33:17, 33:20-33:31, 33:33-33:37, 33:39-33:151, 33:153-33:158, 33:160-33:163, 74:01:01-74:32N, B*15:17:03 |
| 29 | 105 bp 475 bp | *29:19, 29:20, 29:34 *29:82 |  |

4The following HLA-A*29 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| A*29:07, 29:46 | 8 |
| A*29:11, 29:92 | 13 |
| A*29:20, 29:82 | 29 |