HLA-A*03 (101.413-24/04, -24u/04u) 	Lot No: 3L4	Expiry Date: 2024-07-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                              Failed lanes: ____________                             Comments:________________

Gel Picture

PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 2, 3, 17, 18, 47, 51, 55 and 58 may give rise to a lower yield of HLA-specific PCR product than the other A*03 primer mixes.

Primer mixes 17, 19, 31, 50 and 51 may have tendencies of unspecific amplifications.

Primer mixes 23, 28, 42, 57 and 63 have a tendency to giving rise to primer oligomer formation.

Primer mix 9 may give rise to a long unspecific amplification product of approximately 750 bp. This should be disregarded when interpreting the A*03 typings.

Primer mix 64 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1HLA-A*03 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.28.0, April 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-A*03 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product |  | Amplified HLA-A*03 alleles | Other amplified HLA-A alleles |
| --- | --- | --- | --- | --- |
| 6 |  | 190 bp 275 bp | *03:234Q *03:04:01-03:04:03, 03:345:01:01-03:345:01:02 | *01:86, 11:153:01-11:153:02, 11:351Q, 30:34 |
| 10 |  | 65 bp 205 bp | *03:60 *03:02:01-03:02:04, 03:07:01, 03:10, 03:31-03:32, 03:42, 03:73, 03:76, 03:82, 03:90, 03:106, 03:113, 03:133, 03:160, 03:171, 03:198, 03:218, 03:223, 03:236-03:237, 03:242, 03:244, 03:253, 03:274-03:275N, 03:281, 03:285, 03:309, 03:314, 03:318-03:319, 03:332-03:333, 03:348, 03:356, 03:373, 03:385 | *01:12, 01:19, 01:21, 01:126, 01:200, 01:244, 02:156, 02:338, 11:31, 11:35, 11:60, 11:158, 11:183, 11:209, 24:92, 24:458, 30:04:01:01-30:04:03, 30:06, 30:09, 30:17, 30:29, 30:46, 30:77, 30:90, 30:99, 30:103, 30:105, 30:117, 30:139, 30:150, 30:155, 31:03-31:04:01:02, 33:49, 33:144, 68:103:01-68:103:02, 74:23 |
| 13 |  | 55 bp 145 bp | *03:157:01, 03:172, 03:176 *03:01:01:02N | *01:191, 02:45, 11:224, 23:28, 24:77, 31:02:01-31:02:02, 31:07, 31:91, 32:05, 32:31 |
| 14 |  | 105 bp 235 bp | *03:11N *03:59 | *11:137:02N *01:342 |
| 18 |  | 135 bp 215 bp | *03:13 *03:50, 03:64, 03:66, 03:225, 03:266N | *01:98, 11:08, 11:44, 11:183, 11:191, 11:269, 25:43, 30:04:01:01-30:04:03, 30:06, 30:17, 30:29, 30:46, 30:77, 30:90, 30:103, 30:105, 30:117, 30:139, 30:150, 30:155, 36:02, 68:103:01-68:103:02 |
| 19 |  | 125 bp 205 bp 240 bp | *03:202 *03:145:02 *03:15, 03:19, 03:84 |  |
| 20 |  | 60 bp 185 bp 405 bp | *03:61, 03:157:01, 03:172, 03:176 *03:16, 03:79 *03:61 | *01:191, 02:45, 11:224, 23:28, 24:77, 31:02:01-31:02:02, 31:07, 31:91, 32:05, 32:31 |
| 21 |  | 140 bp 195 bp 290 bp | *03:17:01-03:17:02, 03:102, 03:171 *03:74 *03:297N | *01:88:01-01:88:03, 01:236, 02:156, 02:338, 11:119:01-11:119:02, 11:209, 24:51, 24:92, 24:458, 30:61, 30:74, 68:103:01-68:103:02 |
| 22 |  | 90 bp 250 bp | *03:191, 03:236 *03:18, 03:97, 03:122, 03:135, 03:167 | *11:59 *01:01:01:01-01:01:65, 01:01:67-01:02:02, 01:04:01:01N-01:04:01:02N, 01:06-01:19, 01:21-01:33, 01:35-01:62, 01:64-01:65, 01:67:01-01:71, 01:74-01:101, 01:103-01:129, 01:131-01:166, 01:168-01:191, 01:193-01:199, 01:201-01:243, 01:245-01:254, 01:256-01:266:02, 01:268-01:283, 01:285N-01:286, 01:288-01:300, 01:302-01:306, 01:308N-01:314, 01:316-01:343, 11:27, 11:38-11:39, 11:94, 11:209, 11:271, 11:326, 24:51, 24:92, 24:458, 25:59, 30:75, 80:01:01:01-80:06 |
| 23 |  | 95 bp 185 bp 210 bp 240 bp | *03:22:01-03:22:02, 03:248, 03:274 *03:19, 03:103:01-03:103:02, 03:282 *03:56, 03:88, 03:99, 03:177, 03:186 *03:75, 03:347, 03:377 | *31:139, 32:52, C*06:187 *02:237, 02:809, 24:18, 24:204, 24:213, 31:139, 32:52 *11:130, 32:52, 36:02, C*06:187 *24:18, 24:204, 24:213, C*06:187 |
| 27 |  | 95 bp 145 bp 170 bp 240 bp | *03:71, 03:348 *03:25 *03:251 *03:149, 03:245 | *66:37 *11:49 *11:200, 68:193 |
| 28 |  | 170 bp 220 bp 495 bp | *03:26, 03:318 *03:269N *03:41, 03:63, 03:88, 03:382 | *11:77, 11:126, 11:348 *11:130, 36:02, C*06:187 |
| 30 |  | 140 bp 210 bp | *03:27, 03:102 *03:49, 03:139 | *01:157, 11:233 |
| 31 |  | 125 bp 200 bp 230 bp | *03:202 *03:56, 03:88, 03:99, 03:143, 03:177, 03:186, 03:262N *03:28 | *11:130, 32:52, 36:02, C*06:187 |
| 32 |  | 150 bp 220 bp 270 bp | *03:67 *03:29 *03:91N | *02:748N, 32:112N |
| 33 |  | 180 bp 210 bp 285 bp | *03:93 *03:30, 03:152, 03:273 *03:62:01:01-03:62:01:02 | *01:341, 32:37, 80:01:01:01w-80:06w *23:64, 24:104, 32:05, 32:79 *01:144 |
| 34 |  | 125 bp 160 bp | *03:58 *03:31 | *11:273 |
| 35 |  | 235 bp 275 bp | *03:32 *03:55 | *31:21, 74:07, 74:33 |
| 36 |  | 120 bp 160 bp 210 bp | *03:81 *03:33 *03:70 |  |
| 37 |  | 95 bp 120 bp 225 bp | *03:48, 03:168N *03:81 *03:34 | *02:525N |
| 38 |  | 170 bp 195 bp 250 bp | *03:35, 03:130 *03:79, 03:192N *03:73 | *01:167 |
| 39 |  | 200 bp 235 bp | *03:36N, 03:182 *03:45 | *01:109, 11:100, 11:175 *30:77 |
| 40 |  | 240 bp 280 bp 415 bp | *03:37 *03:220 *03:279N |  |
| 41 |  | 185 bp 280 bp | *03:38, 03:192N, 03:223 *03:68N, 03:297N |  |
| 42 |  | 110 bp 190 bp 215 bp 245 bp | *03:52 *03:143 *03:39, 03:238, 03:269N *03:84 |  |
| 43 |  | 150 bp 260 bp | *03:20 *03:47, 03:76 |  |
| 44 |  | 170 bp 240 bp | *03:51:01, 03:130, 03:286N *03:40 |  |
| 45 |  | 150 bp 270 bp | *03:43, 03:186 *03:91N | *30:89, 30:160 *32:112N |
| 46 |  | 115 bp 170 bp 190 bp | *03:44:01-03:44:02, 03:52 *03:213 *03:53 | *02:164:01-02:164:02, 02:380, 29:108, 30:69, 30:109, 31:19, C*01:163 *32:131 *02:575, 29:138, 30:30, 31:69, C*04:341, C*05:123, C*07:35, C*07:352, C*08:48 |
| 48 |  | 70 bp 120 bp 145 bp 190 bp | *03:54 *03:284N *03:46 *03:197N | C*07:86, C*15:198 *02:314N, 32:48N, C*07:55N |
| 49 |  | 70 bp 170 bp 195 bp 545 bp | *03:283N *03:251 *03:161N *03:78 | *11:49 *01:45, 02:453, 02:557, 02:690, 11:108, 23:92, 23:96, 24:271, 33:151, 66:17 |
| 50 |  | 130 bp 165 bp 200 bp | *03:86 *03:213 *03:145:02, 03:262N |  |
| 51 |  | 215 bp 285 bp | *03:50 *03:87 | *11:183, 11:191, 11:269, 25:43, 30:04:01:01-30:04:03, 30:06, 30:17, 30:29, 30:46, 30:77, 30:90, 30:103, 30:105, 30:117, 30:139, 30:150, 30:155, 68:103:01-68:103:02 *01:101, 11:30, 30:92 |
| 52 |  | 90 bp 170 bp 215 bp | *03:96, 03:168N *03:275N *03:89:02, 03:152, 03:172, 03:176, 03:198, 03:273 | *74:14N *01:83:01-01:83:02, 01:191, 02:01:02:01-02:01:02:02, 02:07:14, 02:50, 23:64, 24:104, 32:05, 32:31, 32:79, 74:04, 74:21 |
| 53 |  | 190 bp 225 bp 260 bp 400 bp | *03:234Q, 03:286N *03:150, 03:153 *03:104 *03:242, 03:312 | *33:49 *01:315, 02:45, 02:56:01-02:56:02, 02:78, 02:103, 02:169, 02:195, 02:829, 23:09, 24:129, 24:191, 25:59, 26:19, 26:72, 26:184, 34:19, 66:23, 68:66:01-68:66:02, 68:178 |
| 54 |  | 145 bp 240 bp | *03:195 *03:111, 03:178N |  |
| 56 |  | 180 bp 215 bp | *03:182, 03:275N *03:113, 03:157:02 | *01:109, 11:100, 11:175, 74:14N |
| 57 |  | 180 bp 280 bp 400 bp | *03:134 *03:118, 03:220 *03:112 |  |
| 59 |  | 105 bp 145 bp 240 bp | *03:144 *03:195 *03:128, 03:178N |  |
| 60 |  | 310 bp 400 bp | *03:218 *03:112, 03:129N, 03:279N | *01:246, 11:86 |
| 62 |  | 190 bp 245 bp 375 bp | *03:132, 03:197N *03:204, 03:288 *03:162N | *02:314N, 32:48N, C*07:55N *33:179, C*05:214 |
| 63 |  | 90 bp 135 bp 175 bp 205 bp 525 bp 555 bp | *03:283N *03:284N *03:123:01-03:123:02, 03:171 *03:42, 03:133, 03:139 | *02:113:02N *02:156, 02:338, 11:16, 11:35, 11:57, 11:73, 11:158, 68:103:01-68:103:02 *01:244, 30:09, 31:03-31:04:01:02, 33:49, 33:144, 74:23 *02:113:02N B*56:28Nw |

4The following HLA-A*03 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| A*03:16, 03:61 | 20 | A*03:38, 03:68N | 41 |
| A*03:17:01-03:17:02, 03:74 | 21 | A*03:44:01-03:44:02, 03:53 | 46 |
| A*03:20, 03:47 | 43 | A*03:78, 03:161N | 49 |
| A*03:22:01-03:22:02, 03:103:01-03:103:02, 03:248, 03:282 | 23 | A*03:104, 03:312 | 53 |
| A*03:25, 03:71 | 27 | A*03:118, 03:134 | 57 |
| A*03:27, 03:49 | 30 | A*03:128, 03:144 | 59 |
| A*03:29, 03:67 | 32 | A*03:132, 03:162N | 62 |

5The HLA-A*03 primer set cannot separate the following alleles. These alleles can be distinguished by the HLA-A low resolution kit and/or the HLA-A*74 high resolution kit.

Alleles

A*03:95, 03:215, 03:343, A*74:13

Abbreviations

w: may be weakly amplified

Changes in revision R01 compared to R00:

- Primer mix 53 amplifies the A*03:312 and the A*01:315 alleles. This has been corrected in the Product Insert and the Worksheet.