# Gel Documentation Form and Worksheet

# HLA-A*30 (101.429-12/12u)                                    Lot No: 4L3	Expiry Date: 2024-07-01

# Name:__________________	                     Test Date:_______________	                 Review Date:_______________

Sample ID:______________                            Tested By:_______________              Reviewed By:_______________

DNA Conc.(ng/ul):_________

## PHOTO DOCUMENT

# Interpretation:___________                             Failed lanes: ____________                Comments:________________

Gel Picture

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 4, 12, 14 and 16 may give rise to unspecific amplifications.

Primer mix 3 may give rise to a lower yield of HLA-specific PCR product than the other A*30 primer mixes.

Primer mixes 24 and 25 may have a tendency of giving rise to primer oligomer formation.

In primer mix 19, the positive control band may be weaker than for other HLA-A*30 primer pairs.

Primer mix 32 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1HLA-A*30 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.26.0, October 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3 The following HLA-A*30 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-A*30 alleles | Other amplified Class I alleles |
| --- | --- | --- | --- |
| 5 | 150 bp 245 bp | *30:04:01:01-30:04:03, 30:06, 30:17, 30:29, 30:46, 30:77, 30:90, 30:99, 30:103, 30:105, 30:117, 30:139, 30:150, 30:155 *30:19 | *02:52, 03:82, 24:66, 24:308, 24:333, 68:06 |
| 6 | 80 bp 185 bp 270 bp | *30:06 *30:07 *30:51, 30:106 | *02:121, 02:425, 02:517, 02:850, 03:154:01, 23:47, 24:234, 24:339, 24:347:02, 29:06, 31:51, 33:168, 68:14, 68:158, C*03:125w, C*03:176w |
| 7 | 90 bp 175 bp | *30:28, 30:131 *30:08:01-30:08:02, 30:150 | *26:19 |
| 14 | 210 bp 270 bp | *30:13, 30:16, 30:44, 30:46 *30:51, 30:106 | *02:121, 02:425, 02:517, 02:850, 03:154:01, 23:47, 24:234, 24:339, 24:347:02, 29:06, 31:51, 33:168, 68:14, 68:158, C*03:125w, C*03:176w |
| 16 | 235 bp 265 bp | *30:95 *30:15, 30:33 | *26:187, 36:02 *11:268 |
| 18 | 130 bp 210 bp | *30:32 *30:20, 30:71 |  |
| 19 | 125 bp 200 bp 235 bp | *30:70N *30:31 *30:77 | *03:45 |
| 20 | 210 bp 275 bp | *30:23 *30:34 | *03:04:02-03:04:03, 11:153:01 |
| 21 | 150 bp 185 bp | *30:30, 30:100 *30:24 |  |
| 22 | 180 bp 215 bp | *30:36, 30:73N *30:27N | *03:323N |
| 24 | 105 bp 160 bp 210 bp | *30:78N *30:22 *30:25, 30:57, 30:88, 30:90 | *01:02:01-01:02:02, 01:20, 01:190 |
| 25 | 160 bp 235 bp | *30:40, 30:76N *30:95 | *26:187, 36:02 |
| 26 | 80 bp 190 bp 300 bp | *30:42 *30:121N *30:81 |  |
| 27 | 145 bp 195 bp 465 bp | *30:56 *30:45 *30:132N | *11:166 *01:04:01:01N-01:04:01:02N, 03:21N, 11:21N, 23:07N, 24:11N |
| 29 | 130 bp 190 bp | *30:59N, 30:61, 30:74 *30:73N, 30:121N | *01:57N, 01:88:02-01:88:03, 01:236, 02:156, 02:338, 03:17:01, 03:171, 03:381N, 11:119:01-11:119:02, 11:209, 24:92, 24:458, 68:103:01-68:103:02 |
| 30 | 80 bp 155 bp | *30:37 *30:123N |  |

4The following HLA-A*30 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| A*30:42, 30:81 | 26 |
| A*30:56, 30:132N | 27 |

Abbreviations

w: might be weakly amplified