Gel Documentation Form and Worksheet

HLA-A*31 (101.430-12/12u)	Lot No: 6G7	Expiry Date: 2022-08-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:__________	Failed lanes: ___________ 	   Comments:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 10, 25 and 30 may have tendencies of unspecific amplifications.

Primer mix 14 may give rise to a lower yield of HLA-specific PCR product than the other HLA-A*31 primer mixes

Primer mix 42 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs

1HLA-A*31 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.28.0, September 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-A*31 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-A*31 alleles | Other amplified HLA-A alleles |
| --- | --- | --- | --- |
| 2 | 80 bp 215 bp | *31:67-31:68 *31:02, 31:07-31:08, 31:91, 31:109 | *02:41, 02:80, 02:117, 02:289:01, 02:304, 02:454, 23:45, 24:62, 26:10, 32:28, 32:66, 33:32:01 *02:243:01-02:243:03, 24:82, 29:48, 33:08, 33:53 |
| 5 | 120 bp 285 bp | *31:26, 31:38 *31:03 | *02:660, 03:184, 32:36 *01:143, 11:43, 29:66, 33:13, C*07:449 |
| 6 | 130 bp 165 bp | *31:39, 31:54 *31:05 | *33:07 *23:03:01, 24:21:03, 24:208, 29:03, 29:33, 32:13, 33:10 |
| 8 | 155 bp 220 bp | *31:24, 31:27 *31:09 | *11:01:28, 11:01:77, 29:67, 32:26:02, 33:61, 33:127, 33:142 |
| 10 | 115 bp 160 bp | *31:41 *31:11, 31:56 | *02:24:02, 02:507, 24:21:03, 24:208, 29:33, 29:51, 29:80, 32:02, 32:06, 33:94 |
| 11 | 135 bp 210 bp | *31:26 *31:12, 31:60N | *02:660, 03:184, 32:36 *02:490N, 02:516N, 02:526, 03:269N, 32:89 |
| 14 | 150 bp 220 bp | *31:24 *31:14N | *11:01:28, 11:01:77, 29:67, 32:26:02, 33:61 |
| 15 | 150 bp 225 bp | *31:25 *31:15 | *32:26:01 |
| 16 | 80 bp 165 bp | *31:46 *31:16 | *29:12, 29:92, 33:58 |
| 17 | 160 bp 235 bp | *31:25, 31:40 *31:17 | *32:26:01 |
| 18 | 155 bp 200 bp | *31:43, 31:62 *31:18 | *02:408, 02:590 *23:43, 33:132 |
| 19 | 110 bp 185 bp | *31:19 *31:72 | *02:380, 03:52, 30:109 |
| 21 | 75 bp 155 bp 180 bp 215 bp | *31:35 *31:43, 31:62 *31:21 *31:01:02:03N | *01:07 *02:408, 02:590 *01:07, 02:185, 02:601, 30:55 |
| 22 | 80 bp 155 bp 190 bp | *31:71 *31:27, 31:55 *31:22 | *02:327 *33:127 |
| 23 | 80 bp 165 bp 200 bp | *31:71 *31:40 *31:23 | *02:327 |
| 24 | 150 bp 180 bp 220 bp | *31:55 *31:81 *31:28, 31:89, 31:115 | *02:104 *03:205, 11:43, 33:125, 33:131, 68:29 |
| 25 | 135 bp 165 bp 245 bp | *31:29 *31:56 *31:59 | *02:507, 23:03:01, 23:83, 24:21:03, 24:208, 29:07, 32:89 |
| 26 | 130 bp 185 bp | *31:30, 31:39, 31:97 *31:72 | *02:507, 29:28, 29:79, 32:10, 33:94, B*07:02:40, C*02:02:15, C*04:175 |
| 27 | 85 bp 500 bp | *31:31, 31:67-31:68 | *02:41, 02:80, 02:117, 02:289:01, 02:304, 02:454, 23:45, 24:62, 26:10, 32:28, 32:66, 33:32:01 *02:72, 02:275, 68:156 |
| 28 | 110 bp 175 bp 205 bp | *31:38 *31:32 *31:60N | *02:490N, 02:516N, 03:269N |
| 29 | 115 bp 180 bp 275 bp | *31:48 *31:44, 31:81 *31:36 | *02:140, 26:99, 33:15 *02:104 |
| 30 | 160 bp 185 bp 245 bp | *31:34 *31:87 *31:33 |  |
| 31 | 75 bp 120 bp | *31:37 *31:41, 31:54 | *02:24:02, 02:507, 24:21:03, 24:208, 29:33, 29:51, 29:80, 32:02, 32:06, 32:109, 33:07, 33:94 |

4The A*31:44 and the A*33:15 alleles can be distinguished by the different sizes of the specific PCR products generated by primer mix 29.

The A*31 primer set cannot separate the A*31:89 from the A*33:125 and A*33:131 alleles. These alleles can be distinguished by the HLA-A low resolution and/or HLA-A*33 kits.

5The following HLA-A*31 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| A*31:01:02:03N, 31:35 | 21 | A*31:29, 31:59 | 25 |
| A*31:16, 31:46 | 16 | A*31:36, 31:48 | 29 |

Changes in revision R01 compared to R00:

1.	The expiration date has been altered due to extension of shelf-life.