Gel Documentation Form and Worksheet

HLA-A*32 (101.431-12/12u)	Lot No: 3G4	  Expiry Date: 2021-02-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:__________	Failed lanes: ___________ 	   Comments:_______________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 4 and 20 may have tendencies of unspecific amplifications.

Primer mixes 8 and 10 may give rise to a lower yield of HLA-specific PCR product than the other A*32 primer mixes.

Primer mix 32 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-A*32 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.25.0, July 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-A*32 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-A*32 alleles | Other amplified HLA-A alleles |
| --- | --- | --- | --- |
| 2 | 430 bp 520 bp | *32:01:01:01-32:01:06, 32:01:08-32:01:11, 32:01:13-32:02, 32:05-32:27N, 32:29-32:30:01, 32:31-32:33:01, 32:34-32:65, 32:67-32:77, 32:79-32:83, 32:85-32:93, 32:95-32:100, 32:102-32:110 *32:01:01:01-32:01:05, 32:01:07-32:01:29, 32:04-32:09, 32:11Q-32:21, 32:23-32:77, 32:79-32:83, 32:85-32:100, 32:102-32:110 | *29:13 *02:81, 02:124, 23:36, 25:19:01-25:19:02, 25:30, 29:13 |
| 6 | 120 bp 520 bp | *32:20 *32:04, 32:52 | *03:152, 03:219, 24:18, 24:204, 24:213 |
| 8 | 165 bp 215 bp | *32:19N *32:07 | *01:02w, 01:20w, 03:72, 11:88, 23:09w, 23:51, 24:24, 24:67, 24:145, 24:156, 24:191, 24:290, 24:392, 26:16, 29:37, 29:56, 30:01:01-30:04:02, 30:06, 30:09-30:20, 30:23-30:30, 30:32-30:54, 30:56-30:59N, 30:61-30:78N, 30:80-30:129, 33:119, 68:45, 68:117 |
| 11 | 120 bp 165 bp | *32:20 *32:09 |  |
| 12 | 130 bp 195 bp | *32:10 *32:16 | *02:507, 29:28, 29:79, 31:30, 31:97, 33:94, B*07:02:40, C*02:02:15, C*04:175 |
| 13 | 155 bp 200 bp | *32:11Q *32:15 | *01:51, 02:55, 02:644, 03:24, 25:03, 25:30, 26:20, 34:08, 68:71 |
| 15 | 165 bp 220 bp | *32:13 *32:18 | *23:03:01, 24:21:03, 24:208, 29:03, 29:33, 31:05, 33:10 |

| 17 | 75 bp 200 bp | *32:21 *32:44 | *29:62, 31:45, 33:16 *33:75 |
| --- | --- | --- | --- |
| 18 | 130 bp 185 bp | *32:24 *32:60 |  |
| 19 | 110 bp 185 bp | *32:25, 32:45N *32:60 |  |
| 20 | 125 bp 220 bp | *32:23 *32:54 | *33:46 *02:294, 34:01:01?-34:01:02?, 34:05?, 66:08 |
| 22 | 175 bp 230 bp | *32:56N *32:27N |  |
| 23 | 80 bp 225 bp | *32:28, 32:66 *32:53 | *02:41, 02:80, 02:117, 02:289:01, 02:304, 02:454, 23:45, 24:62, 26:10, 31:67-31:68, 33:32:01 *02:480, 33:39, 68:176 |

4The following HLA-A*32 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| A*32:10, 32:16 | 12 |
| A*32:21, 32:44 | 17 |
| A*32:23, 32:54 | 20 |
| A*32:28, 32:53, 32:66 | 23 |

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.