PAGE  
Page 1(4)

July 2007
HLA-A*68
Rev. No: 00
X65



	PHOTO DOCUMENT


Olerup SSP AB
Olerup SSPTM HLA-A*68
Prod. No:
101.418-12/04 – including Taq pol.

101.418-12u/04u – without Taq pol.
Lot No:    
X65
Expiry Date:    
2009-May-01
Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 32 primer mixes used for HLA-A*68 SSP typing.
	Primer Mix
	Size of spec. PCR product1
	Size of control band2
	Amplified HLA-A*68 alleles
	Other amplified HLA-A alleles3

	  14
	195 bp
	  800 bp
	680101-68020102, 6806-6814, 6816-6819, 6821-6830, 6832-6835, 6837, 6838
	0113, 0234-023502, 0256, 0262, 0278, 9203, 03010101-0307, 0309-0329, 110101-110107, 1103-1113, 1115, 1120-1130, 2419, 3018, 3401-3408, 6601, 6602, 6604, 6606, 6901

	  2
	435 bp
	1070 bp
	680101-680106, 680301-6805, 6808, 6809, 6811N, 6812, 6816, 6817, 6819-6826, 6829, 6832, 6833, 6835-6838
	0324, 1110, 3402-3404, 3407, 3408

	  3
	195 bp
	1070 bp
	68020101-68020102, 6815, 6818N, 6828, 6831, 6834
	

	  45
	180, 220 bp
	1070 bp
	680301-680302, 6805, 6820, 6830
	

	  56
	140, 235 bp
	  800 bp
	6804, 6825
	

	  6
	225 bp
	1070 bp
	6805, 6815, 6820
	

	  77
	135, 155 bp
	1070 bp
	6806, 6832
	0252

	  8
	160 bp
	1070 bp
	6807
	

	
	
	
	
	

	  9
	160 bp
	1070 bp
	6808, 6818N
	030102, 1124, 8001

	10
	165 bp
	  800 bp
	6809, 6826
	0112, 0119, 0310, 110101-1107, 1109weakly, 1110-1121N, 1126-1130, 2417, 2441, 2905

	118
	160, 200 bp
	  800 bp
	6810, 6814, 6817
	

	12
	405 bp
	1070 bp
	6811N
	

	13
	215 bp
	1070 bp
	6812
	3601-3604faintly

	14
	195 bp
	1070 bp
	6813
	

	159,10
	145, 205 bp
	1070 bp
	6812?, 6814, 6819
	0317, 2418

	1611
	150, 215 bp
	1070 bp
	6815, 6816, 6831
	

	
	
	
	
	

	1712
	165, 185 bp
	1070 bp
	68020101-68020102, 6815, 6818N, 6820, 6827, 6828, 6831, 6834
	0202weakly, 020301-020302, 0205weakly, 0208weakly, 0222, 0247weakly, 0249, 0250, 0263weakly, 9202weakly, 9204, 9215weakly, 2302, 2406, 250101-2503, 2505, 2506, 260101-260104, 2602-260702, 2609-2626, 2628-2634, 3210, 3401, 3405, 4301, 6601-6605

	1813
	195, 225 bp
	1070 bp
	6821, 6826
	01010101-0104, 0106-0125, 0304, 0318, 1127, 2417, 2441, 2903, 8001

	199
	325 bp
	1070 bp
	6822
	

	209,14
	  90 bp
	1070 bp
	6823
	0120, 9214

	2115
	165, 215 bp
	1070 bp
	6824, 6828
	0249

	224
	165 bp
	  800 bp
	680101-680106, 680301-6805, 6807, 6810-6814, 6816, 6817, 6819-6825, 6829, 6830, 6832, 6835-6838
	01010101-0104Nweakly, 0107-0111Nweakly, 0113-0118Nweakly, 0120-0124weakly, 0318weakly, 1108, 3601-3604weakly

	23
	185 bp
	1070 bp
	680101-680106, 680301-6805, 6807-6814, 6816, 6817, 6819-6825, 6829, 6830, 6832, 6832, 6835-6838
	9214, 03010101-0317, 0320-0329, 1104, 1127, 29010101-290201, 290203, 2904-2916, 3103, 3104, 3106, 3204, 3402-3404, 3407, 3408, 3601-3603

	2416
	375 bp
	1070 bp
	
	0255, 2622, 6901

	
	
	
	
	

	25
	240 bp
	1070 bp
	6829
	

	26
	185 bp
	1070 bp
	6833
	

	27
	190 bp
	1070 bp
	6834
	2604

	28
	200 bp
	1070 bp
	6835
	

	29
	255 bp
	  800 bp
	6836
	

	3017
	145, 245 bp
	1070 bp
	6837
	250101-2506, 260101-2610, 2612, 2613, 2615-2617, 2619-2627, 2629-2634, 3401, 3406, 4301, 6601-6606

	31
	205 bp
	1070 bp
	6838
	

	32
	225 bp
	1070 bp
	6821
	0304


1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-A*68 SSP subtypings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 bp or more. Size differences shorter than 20 bp are not given. For high resolution SSP kits the length of the specific PCR product(s) of the alleles amplified by these primer mixes are given. 

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.  

PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most tubes, or a band of 800 base pairs, for some tubes.

Tube number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-A*68 subtyping.

In addition, tubes number 5, 10, 11, 22 and 29 contain the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases. 

3The nucleotide sequences of the A*68020103 and A*6839 alleles are not yet retrievable.

Due to the sharing of sequence motifs between HLA-A alleles non-HLA-A*68 alleles will be amplified by primer mixes 1, 2, 7, 9, 10, 13, 15, 17, 18, 21, 20, 22, 23, 24, 27 and 30.

4Primer mixes 1 and 22 may yield less specific PCR product than the other A*68 primer mixes.

5Primer mix 4: Specific PCR fragment of 180 bp in the A*6830 allele. Specific PCR fragment of 220 bp in the A*680301-680302, A*6805 and A*6820 alleles.

6Primer mix 5: Specific PCR fragment of 140 bp in the A*6825 allele. Specific PCR fragment of 235 bp in the A*6804 allele.

7Primer mix 7: Specific PCR fragment of 135 bp in the A*6806 and A*0252 alleles. Specific PCR fragment of 155 bp in the A*6832 allele.

8Primer mix 11: Specific PCR fragment of 160 bp in the A*6817 allele. Specific PCR fragment of 200 bp in the A*6810 and A*6814 alleles.

9Primer mixes 15, 19 and 20 have a tendency of giving rise to a primer oligomer artifact.

10Primer mix 15: Specific PCR fragment of 145 bp in the A*6819 and A*0317, A*2418 alleles. Specific PCR fragment of 205 bp in the A*6814 allele.

11Primer mix 16: Specific PCR fragment of 150 bp in the A*6816 allele. Specific PCR fragment of 215 bp in the A*6815 and A*6831 alleles.

12Primer mix 17: Specific PCR fragment of 165 bp in the A*68020101-68020102, A*6815, A*6818N, A*6827, A*6828, A*6831, A*6834 and A*0202, A*020301-020302, A*0205, A*0208, A*0222, A*0247, A*0249, A*0250, A*0263, A*9202, A*9204, A*9215, A*2302, A*2406, A*250101-2503, A*2505, A*2506, A*260101-260104, A*2602-260702, A*2609-2626, A*2628-2634, A*3210, A*3401, A*3405, A*4301, A*6601 to A*6605 alleles. Specific PCR fragment of 185 bp in the A*6820 allele.

13Primer mix 18: Specific PCR fragment of 195 bp in the A*6826 and A*01010101-0104N, A*0106-0125, A*0318, A*1127A*2417, A*2441, A*2903 and A*8001 alleles. Specific PCR fragment of 225 bp in the A*6821 and A*0304 alleles.

14Short specific PCR fragments are less intense and not as sharp as longer specific bands.

15Primer mix 21: Specific PCR fragment of 165 bp in the A*6828 and A*0249 alleles. Specific PCR fragment of 215 bp in the A*6824 allele.

16Primer mix 24 may give to non-specific amplifications.

17Primer mix 30: Specific PCR fragment of 145 bp in the A*250101-2506, A*260101-2610, A*2612, A*2613, A*2615-2617, A*2619-2627, A*2629-2634, A*3401, A*3406, A*4301, A*6601-6606 alleles. Specific PCR fragment of 245 bp in the A*6837 allele.

‘?’, nucleotide sequence information not available for the primer matching sequence.
[image: image1.emf]Tube No. 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24 25 26 27 28 29 30 31 32

HLA-A allele

1

*680101-680106 1 2 22 23

*68020101-68020101 1 3 17

*680301-680302 2 4 22 23

*6804 2 5 22 23

*6805 2 4 6 22 23

*6806 1 7

*6807 1 8 22 23

*6808 1 2 9 23

*6809 1 2 10 23

*6810 1 11 22 23

*6811N 1 2 12 22 23

*6812 1 2 13 ? 22 23

*6813 1 14 22 23

*6814 1 11 15 22 23

*6815 3 6 16 17

*6816 1 2 16 22 23

*6817 1 2 11 22 23

*6818N 1 3 9 17

*6819 1 2 15 22 23

*6820 2 4 6 17 22 23

*6821 1 2 18 22 23 32

*6822 1 2 19 22 23

*6823 1 2 20 22 23

*6824 1 2 21 22 23

*6825 1 2 5 22 23

*6826 1 2 10 18

*6827 1 17

*6828 1 3 17 21

*6829 1 2 22 23 25

*6830 1 4 22 23

*6831 3 16 17

*6832 1 2 7 22 23

*6833 1 2 22 23 26

*6834 1 3 17 27

*6835 1 2 22 23 28

*6836 2 22 23 29

*6837 1 2 22 23 30

*6838 1 2 22 23 31

*01010101-0104N, 0107-0111N, 0114-0118N, 

0121-0124

18 w

*0106, 01252903 18

*0112, 0119 10 18

*0113, 0318 1 18 w

*0120 18 20 w

*0202, 0205, 0208, 0247, 0263, 9202, 9215 w

*020301-020302, 0222, 0250, 9204, 2302, 2406, 

2611N, 2614, 2618, 2628, 3210

17

*0234-023502, 0256, 0262, 0278, 9203, 0319, 

1122, 1123, 1125, 2419, 3018, 3406, 6606

1

*0249 17 21

*0252 7

*0255 24

*9214 20 23

*03010101-03010103, 030103-0303N, 0305-

0307, 0309, 0311N-0316, 0320-0323, 0325-0329

1 23

*030102 1 9 23

*0304 1 18 23 32

*0308, 29010101-290201, 290203, 2904, 2906-

2916, 3103, 3104, 3106, 3204

23


1The nucleotide sequences of the A*68020103 and A*6839 alleles are not yet retrievable.
2Primer mix 4: Specific PCR fragment of 180 bp in the A*6830 allele. Specific PCR fragment of 220 bp in the A*680301-680302, A*6805 and A*6820 alleles.

Primer mix 5: Specific PCR fragment of 140 bp in the A*6825 allele. Specific PCR fragment of 235 bp in the A*6804 allele.

Primer mix 7: Specific PCR fragment of 135 bp in the A*6806 and A*0252 alleles. Specific PCR fragment of 155 bp in the A*6832 allele.

Primer mix 11: Specific PCR fragment of 160 bp in the A*6817 allele. Specific PCR fragment of 200 bp in the A*6810 and A*6814 alleles.

Primer mix 15: Specific PCR fragment of 145 bp in the A*6819 and A*0317, A*2418 alleles. Specific PCR fragment of 205 bp in the A*6814 allele.

Primer mix 16: Specific PCR fragment of 150 bp in the A*6816 allele. Specific PCR fragment of 215 bp in the A*6815 and A*6831 alleles.

Primer mix 17: Specific PCR fragment of 165 bp in the A*68020101-68020102, A*6815, A*6818N, A*6827, A*6828, A*6831, A*6834 and A*0202, A*020301-020302, A*0205, A*0208, A*0222, A*0247, A*0249, A*0250, A*0263, A*9202, A*9204, A*9215, A*2302, A*2406, A*250101-2503, A*2505, A*2506, A*260101-260104, A*2602-260702, A*2609-2626, A*2628-2634, A*3210, A*3401, A*3405, A*4301, A*6601 to A*6605 alleles. Specific PCR fragment of 185 bp in the A*6820 allele.

Primer mix 18: Specific PCR fragment of 195 bp in the A*6826 and A*01010101-0104N, A*0106-0125, A*0318, A*1127A*2417, A*2441, A*2903 and A*8001 alleles. Specific PCR fragment of 225 bp in the A*6821 and A*0304 alleles.

Primer mix 21: Specific PCR fragment of 165 bp in the A*6828 and A*0249 alleles. Specific PCR fragment of 215 bp in the A*6824 allele.

Primer mix 30: Specific PCR fragment of 145 bp in the A*250101-2506, A*260101-2610, A*2612, A*2613, A*2615-2617, A*2619-2627, A*2629-2634, A*3401, A*3406, A*4301, A*6601-6606 alleles. Specific PCR fragment of 245 bp in the A*6837 allele.

‘w, may be weakly amplified.

‘f’, may be faintly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

[image: image2.emf]*0310, 1104 1 10 23

*0317 1 15 23

*0324 1 2 23

*110101-110107, 1103, 1105-1107, 1111-1113, 

1115, 1120, 1121N, 1126, 1128-1130

1 10

*110201-110202, 1114, 1116-1119 10

*1108 1 22

*1109 1 w

*1110 1 2 10

*1124 1 9

*1127 1 10 18 23

*2417, 2441 10 18

*2418 15

*250101-2503, 2505, 2506, 260101-260104, 

2602, 2603, 2605-260702, 2609, 2610, 2612, 

2613, 2615-2617, 2619-2621, 2623-2626, 2629-

2634, 4301, 6603, 6605

17 30

2504, 260105, 2608, 2627 30

*2604 17 27 30

*2622 17 24 30

*2905 10 23

*3401, 3405, 6601, 6602, 6604 1 17 30

*3402-3404, 3407, 3408 1 2 23

*3601-3603 f w 23

*3604 f w

*6901 1 24

*8001 9 18

Length of spec. PCR product(s)

2

195 435 195 180 140 225 135 160 160 165 160 405 215 195 145 150 165 195 325

90

165 165 185 375 240 185 190 200 255 145 205 225

220 235 155 200 205 215 185 225 215 245

Tube No. 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24 25 26 27 28 29 30 31 32

Specific Reaction

1st match: A*

2nd match: A*

Control Reaction

Failed Controls

False Positive

False Negative








Olerup SSP AB
In U.S. “For Research Use Only.
www.olerup.com
Hasselstigen 1
Not For Use in Diagnostic Procedures.”

SE-133 33 Saltsjöbaden

Sweden

