Gel Documentation Form and Worksheet

HLA-B*14 (101.524-12/12u) 	       Lot No: 7G2	Expiry Date: 2021-03-01

Name:___________________

Sample ID:_______________

DNA Conc.(ng/ul):_________

Test Date: ____________

Tested By: _____________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________	Failed lanes: ___________ 	 Comments:_________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 10 may give rise to a lower yield of HLA-specific PCR product than the other B*14 primer mixes.

Primer mix 4, 13 and 14 may have tendencies of unspecific amplifications.

Primer mix 22 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pair.

1HLA-B*14 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.27.0, January 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-B*14 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-B*14 alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- |
| 4 | 85 bp 270 bp | *14:03 *14:20 | *07:218, 07:220, 18:13, 18:30, 51:29, 51:82, 51:217, 52:24, C*01:30, C*08:51, C*08:114, C*14:76 |
| 5 | 135 bp 220 bp | *14:07N *14:04 | *39:40:01N-39:40:02N |
| 6 | 100 bp 175 bp 530 bp | *14:05, 14:29 *14:17 *14:29 | *07:115, 07:246, 08:69, 38:19, 39:03:01:01-39:03:01:02, 39:24:01-39:24:02, 39:37, 39:76, 39:120, 40:314, 41:18, 42:06, 42:20, 48:07, C*08:08:01-08:08:02 |
| 9 | 160 bp 280 bp | *14:09 *14:32 | C*05:16, C*05:85, C*05:107, C*06:129, C*07:364, C*08:12 |
| 10 | 110 bp 180 bp | *14:29 *14:11, 14:28 | *08:01:24, 15:189, 35:03:17, 35:26, 38:33, 39:04, 49:26 |
| 15 | 100 bp 270 bp | *14:15 *14:10 | *55:77 *18:22, 18:69, 18:105, 35:21, 35:24:01-35:24:02, 35:81, 35:96, 35:109, 35:157, 35:188, 35:190, 35:233, 35:287, 37:04:01-37:04:02, 40:28, 44:190, 44:203, 44:275, 51:04, 51:46, 51:56:01-51:56:03, 51:139, 51:227, 53:02, 53:06, 53:28, 57:14:01-57:14:02, 58:09, 58:76, C*03:384, C*15:39, C*15:130 |
| 16 | 90 bp 175 bp 280 bp | *14:23 *14:16 *14:32 | *07:239, C*01:101, C*01:113, C*15:55 C*05:16, C*05:85, C*05:107, C*06:129, C*07:364, C*08:12 |

4The B*14:08:01-14:08:02 and the B*39:43 alleles will give rise to identical amplification patterns with the HLA-B*14 subtyping kit. These alleles can be distinguished by the HLA-B low resolution kit and/or the HLA-B*39 subtyping kit.

5The following HLA-B*14 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| B*14:16, 14:23 | 16 |

‘w’, might be weakly amplified.