Gel Documentation Form and Worksheet

HLA-B*27 (101.521-24/04, -24u/04u) 	Lot No: 8G7	Expiry Date: 2021-04-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                           Failed lanes: ___________ 	                           Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 28, 30, 31, 32 and 39 may have tendencies of unspecific amplifications.

Primer mix 43 has a tendency to giving rise to primer oligomer formation.

Primer mix 30 may give rise to a long unspecific amplification product. This band should be disregarded when interpreting the typing results.

Primer mix 47 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-B*27 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.28.0, April 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-B*27 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-B*27 alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- |
| 10 | 90 bp 185 bp | *27:06, 27:35, 27:41, 27:78, 27:91, 27:106-27:107, 27:154 *27:60 | *15:129, 15:395, 18:02, 38:22, C*07:05, C*07:538 |
| 12 | 210 bp 310 bp 425 bp | *27:74 *27:43 *27:13 | *08:100, 15:292, 15:401, 35:113, 39:62, 40:202, 44:98, 46:27, 52:31:01-52:31:02 |
| 13 | 215 bp 370 bp 400 bp | *27:47 *27:12:01:01-27:12:01:02, 27:16, 27:92, 27:119, 27:153 *27:53 | *18:33, 37:57, 40:75?, 40:322? |
| 14 | 120 bp 345 bp | *27:34 *27:17 | *35:01:30, 40:02:07, 40:06:02, 51:01:52, 57:01:16, 57:03:02, C*02:06:02, C*02:23, C*02:36:01, C*02:68 |
| 17 | 100 bp 310 bp | *27:14, 27:81 *27:51 | *44:20, 44:47, 44:100, 44:197, 50:20, 51:85, 51:192, 51:214, 52:26 |
| 18 | 175 bp 415 bp | *27:24 *27:49 | *07:174, 07:202, 07:222, 07:279, 40:31, 40:45, 40:80, 48:01:01:01-48:01:06, 48:04:01-48:04:02, 48:06-48:07, 48:09, 48:11, 48:15-48:16, 48:18-48:20, 48:22, 48:24, 48:27-48:39, 48:41-48:46, 81:01:01-81:08 |
| 19 | 95 bp 240 bp | *27:35 *27:25, 27:38, 27:47, 27:50:01-27:50:02, 27:63, 27:76, 27:139 | *38:22 *15:129, 15:342, 15:395 |
| 20 | 95 bp 175 bp 215 bp 345 bp | *27:09 *27:36, 27:173 *27:25, 27:45, 27:108, 27:115, 27:161 *27:69 | *48:21, 48:26 *07:197, 15:129, 15:342 *15:227, 37:23, 40:213, 51:32, 58:52 |
| 21 | 110 bp 160 bp | *27:04:01-27:04:06, 27:06, 27:24-27:25, 27:66N, 27:68-27:69, 27:100, 27:103, 27:105, 27:109, 27:112-27:115, 27:138, 27:147, 27:149, 27:165, 27:173 *27:39, 27:137 | *37:47 |
| 22 | 165 bp 195 bp 235 bp 375 bp | *27:55, 27:149 *27:37 *27:15, 27:28, 27:62, 27:71 *27:26, 27:31, 27:101, 27:157 | *18:43, 40:210?, 40:229 *18:02, A*32:89, C*07:05, C*07:538 *40:166?, 82:01-82:03 |
| 26 | 105 bp 145 bp 235 bp | *27:52 *27:64N *27:72 | *37:60 *13:103N, 40:118N, 40:291N |
| 27 | 215 bp 320 bp | *27:104 *27:57 | *37:27 |
| 28 | 145 bp 325 bp | *27:64N *27:48 | *13:103N, 40:118N, 40:291N |
| 29 | 70 bp 105 bp 210 bp | *27:54, 27:123 *27:67 *27:74 |  |
| 30 | 160 bp 240 bp | *27:59N, 27:137 *27:72-27:73 |  |
| 31 | 125 bp 160 bp 265 bp | *27:65N *27:56 *27:118 | A*24:185N |
| 32 | 90 bp 370 bp | *27:32, 27:70 *27:85 | *07:28, 07:118, 08:14, 08:87, 15:17:01:01-15:17:07, 15:162, 15:168, 15:177, 15:196,15:208, 15:216, 15:230, 15:233, 15:269, 15:273, 15:339, 15:356, 15:396, 15:403, 15:411, 15:423-15:424, 15:446, 35:234, 40:124:02, 40:180, 40:240, 44:02:33, 44:41:01, 44:54, 44:65, 44:106, 44:135, C*02:07, C*02:97 |
| 34 | 160 bp 265 bp | *27:82 *27:118 |  |
| 35 | 120 bp 280 bp | *27:134 *27:83 |  |
| 36 | 190 bp 350 bp | *27:146 *27:94N | *37:33N |
| 41 | 105 bp 155 bp | *27:67 *27:91 |  |
| 44 | 225 bp 255 bp 290 bp | *27:66N *27:01, 27:142 *27:58 | *13:23, 13:55, 18:09, 37:10, 40:47, 40:96, 40:157, 40:201, 44:15:01:01-44:15:01:02, 44:55, 44:103, 44:188, 44:212, 49:02 |

4The following HLA-B*27 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| B*27:37, 27:55, 27:101 | 22 |
| B*27:53, 27:119 | 13 |
| B*27:69, 27:173 | 20 |
| B*27:83, 27:134 | 35 |

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.