Gel Documentation Form and Worksheet

HLA-B*40 (101.523-24/03, -24u/03u) 	Lot No: 5F8	Expiry Date: 2020-04-01

Name:__________________

Sample ID:_______________

DNA Conc.(ng/ul):_________

Test Date: ___________

Tested By: ___________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 20, 21, 23, 34, 38, 39, 41, 45, 54, 72, 74 and 91 may have tendencies of unspecific amplifications.

Primer mix 7 may faintly amplify the C*17 alleles.

Primer mixes 14, 27, 35, 45, 55, 56, 65, 69 and 70 have a tendency to giving rise to primer oligomer formation.

Primer mixes 46, 48 and 70 may give rise to a lower yield of HLA-specific PCR product than the other B*40 primer mixes, most pronounced in primer mix 46.

Primer mix 95 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

s

1HLA-B*40 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.28.0, April 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-B*40 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-B*40 alleles | Other amplified HLA class I alleles |
| --- | --- | --- | --- |
| 16 | 75 bp 160 bp | *40:15 *40:01:46, 40:14:01, 40:72:01, 40:77, 40:121, 40:136, 40:237 | *07:02:13, 07:02:35, 07:68:02, 07:151:02, 07:209, 07:262, 18:35, 18:110, 35:66, 42:05:01-42:05:02, 48:12, C*07:38:01-07:38:02 |
| 18 | 105 bp 175 bp | *40:01:06, 40:26 *40:197 | *07:02:11, 07:05:03, 15:01:02, 15:09:01-15:09:02, 15:243, 15:252, 15:277-15:278, 15:361, 35:251, 51:06:01, 51:06:03, 51:59, 51:64, 51:135, 51:194, C*02:02:23 *18:38 |
| 20 | 135 bp 285 bp | *40:14:01-40:16, 40:23, 40:32, 40:45, 40:98, 40:114:01-40:114:02, 40:121, 40:136, 40:158, 40:231, 40:355 *40:15-40:16, 40:31-40:32, 40:45, 40:80, 40:98 | *07:02:01:01-07:18:02, 07:20-07:32, 07:34-07:39, 07:41-07:52, 07:54-07:59, 07:61-07:99, 07:101-07:105, 07:107-07:121, 07:123-07:138, 07:140-07:146, 07:148-07:150, 07:152-07:164, 07:166-07:201N, 07:203-07:234, 07:236-07:278, 07:280-07:294, 08:79, 08:156, 15:138, 15:230, 15:241, 15:297, 18:35, 18:110, 35:66, 37:07, 41:08, 42:05:01-42:05:02, 44:150, 48:05, 48:08, 48:10, 48:12, C*02:23, C*03:260, C*04:77 *07:02:01:01-07:02:38, 07:02:40-07:06:02, 07:08:01-07:11, 07:13, 07:15-07:17, 07:20-07:32, 07:34-07:39, 07:41-07:52, 07:54-07:59, 07:61-07:76, 07:79-07:83, 07:85-07:99, 07:101-07:121, 07:123-07:136:02, 07:138, 07:140-07:161N, 07:164, 07:166-07:192, 07:194-07:195, 07:199-07:213, 07:215-07:224, 07:226, 07:228:01-07:234, 07:236-07:238, 07:240-07:247, 07:249-07:254, 07:256-07:294, 08:79, 08:156, 35:66, 37:07, 42:05:01-42:05:02, 48:01:01-48:01:06, 48:05-48:12, 48:14-48:16, 48:18-48:20, 48:22, 48:27-48:39, 48:41, 81:01-81:08 |
| 23 | 95 bp 160 bp 205 bp | *40:29 *40:294 *40:264 | *08:144, 18:77, 39:66, C*03:104, C*07:581, C*08:122, C*12:25 |
| 34 | 190 bp 235 bp | *40:118N *40:42, 40:60, 40:124:01-40:124:02, 40:172, 40:280 | *07:14, 07:17, 07:28, 07:96:01-07:96:02, 07:115, 07:197, 07:225, 07:260, 07:284, 15:230, 15:297, 27:36, 44:150, 48:17, 48:26, C*03:260, C*04:77, C*07:38:01-07:38:02 |
| 36 | 90 bp 135 bp | *40:273 *40:10:01:01-40:10:02, 40:40, 40:49, 40:58, 40:110, 40:129, 40:164, 40:210, 40:219, 40:229, 40:325, 40:342 | *15:46, 15:53, 15:106, 18:107, 35:47, 35:154, 41:33, 41:47, 44:02:01:01-44:02:31, 44:02:33-44:03:15, 44:03:17-44:03:29, 44:03:31-44:05:04, 44:09-44:14, 44:16-44:17, 44:19N-44:39, 44:41:01-44:43:02, 44:45-44:48, 44:50:01-44:54, 44:56N, 44:58N-44:59:02, 44:61N-44:80, 44:82, 44:84:01-44:102, 44:104-44:110, 44:112-44:129, 44:132-44:133, 44:135-44:144, 44:146-44:155, 44:157-44:187, 44:189-44:193, 44:195N-44:199, 44:201-44:209, 44:211, 44:213-44:224, 44:226-44:245, 44:247-44:269, 47:01:01:01-47:10, 49:10, 49:24, 49:31, 50:13, 50:43, 53:30, 57:45, 57:69, C*16:85 |
| 37 | 125 bp 270 bp | *40:51, 40:113, 40:163 *40:28 | *14:10, 18:22, 18:69w, 18:105, 35:21, 35:24:01-35:24:02, 35:81, 35:96, 35:109, 35:157, 35:188, 35:190, 35:233, 35:287, 37:04:01w-37:04:02w, 44:190, 44:203, 51:04, 51:46, 51:56:01-51:56:03, 51:139, 53:02, 53:06, 53:28, 57:14:01-57:14:02, 58:09, 58:76, C*15:39, C*15:130 |
| 40 | 120 bp 150 bp | *40:38, 40:52, 40:59-40:60, 40:222, 40:352 *40:216N | *15:116, 15:124, 15:352, 15:412, 35:331, 46:06, 46:43, 48:23, 57:59 |
| 44 | 110 bp 210 bp | *40:54, 40:242, 40:263N | * 51:174:02 |
| 45 | 175 bp 235 bp 260 bp | *40:56 *40:55 *40:301 |  |
| 48 | 215 bp 245 bp | *40:61 *40:73 | *18:06, 46:14 |
| 53 | 165 bp 245 bp | *40:65 *40:171 | *15:193, 39:53, C*04:155, C*07:107, C*07:224 |
| 56 | 100 bp 165 bp 270 bp | *40:69 *40:97 *40:221 |  |
| 59 | 85 bp 115 bp 150 bp 190 bp | *40:127 *40:265N *40:331, 40:346 *40:81 | *44:143, 44:261 *15:175 |
| 60 | 135 bp 240 bp | *40:82 *40:88, 40:225 | *07:196, 27:80 |
| 61 | 80 bp 135 bp | *40:02:04, 40:10:02, 40:83, 40:248 *40:234 | *07:133, 15:46, 15:106, 15:143, 35:19, 35:47, 35:154, 35:274, 37:06:01-37:06:02, 44:84:01, 44:194, 44:205:01-44:205:02, 44:225, 44:269, 48:41, 51:112, 53:39, 57:45 |
| 62 | 150 bp 195 bp 245 bp | *40:182, 40:205 *40:84 *40:119 |  |
| 63 | 155 bp 180 bp 240 bp | *40:255 *40:85 *40:133Q |  |
| 65 | 130 bp 170 bp | *40:02:01:01-40:02:17, 40:02:19-40:02:25, 40:04-40:06:16, 40:08, 40:11:01-40:11:02, 40:13-40:16, 40:19, 40:26-40:29, 40:35:01-40:35:02, 40:37, 40:39-40:40, 40:50, 40:53, 40:56-40:58, 40:64, 40:68, 40:70:01-40:71, 40:75, 40:78, 40:82-40:83, 40:86, 40:90, 40:93-40:97, 40:99, 40:103-40:105, 40:107, 40:109-40:111, 40:115, 40:119, 40:121-40:122, 40:127, 40:129, 40:133Q, 40:136, 40:142N-40:145, 40:148, 40:157-40:158, 40:161-40:162, 40:164-40:165, 40:167, 40:169, 40:173-40:174, 40:176-40:177, 40:181, 40:184, 40:189-40:190, 40:200-40:203, 40:205-40:206, 40:209, 40:211, 40:214, 40:219-40:220, 40:224-40:226, 40:229-40:230, 40:232, 40:237, 40:243-40:244, 40:246, 40:248, 40:254-40:256N, 40:266, 40:268-40:271, 40:274-40:276, 40:283, 40:287, 40:289, 40:290w, 40:291N-40:297, 40:302-40:304, 40:305w, 40:306, 40:309, 40:311, 40:313-40:314, 40:317-40:318, 40:320, 40:322, 40:331, 40:334-40:335, 40:337N, 40:340, 40:342-40:343, 40:345N, 40:347-40:351, 40:354 *40:15-40:16, 40:23, 40:128, 40:161, 40:198, 40:324 | *07:05:01:01-07:06:02, 07:32, 07:34, 07:40, 07:53, 07:69, 07:78, 07:80, 07:97, 07:105, 07:112, 07:123, 07:137-07:138, 07:140, 07:176, 07:182N, 07:206-07:207, 07:209-07:210, 07:213-07:214, 07:249, 07:258, 07:262, 07:264, 07:269-07:270, 07:278, 07:283, 07:287, 07:293, 08:01:01:01-08:01:11, 08:01:13-08:05, 08:08N-08:12:01, 08:12:03-08:13, 08:15-08:27, 08:29-08:36, 08:38-08:48, 08:50-08:63, 08:65-08:73, 08:75-08:82N, 08:84w, 08:85-08:86N, 08:88, 08:90-08:106, 08:108w, 08:109-08:110, 08:112-08:155, 08:157-08:180, 13:03, 13:22:01-13:22:02, 13:40w, 13:45, 13:48, 13:62, 13:86w, 14:13, 14:53, 15:09:01-15:09:02, 15:30, 15:42, 15:48, 15:58, 15:73, 15:83, 15:86, 15:150, 15:188, 15:224, 15:229, 15:243, 15:252, 15:303, 15:324, 15:329, 15:361, 18:26, 18:61, 18:97, 27:07:01-27:07:05, 27:11, 27:20, 27:27, 27:33, 27:43, 27:130, 27:153, 35:02:01:01w, 35:02:01:02, 35:02:02w-35:02:08w, 35:04:01-35:04:03, 35:06, 35:09:01-35:09:03, 35:12:01-35:12:04, 35:18, 35:22, 35:31, 35:44, 35:59:01-35:60, 35:81w, 35:83w, 35:87-35:88, 35:95w, 35:129Nw, 35:149, 35:154w, 35:157w, 35:162w, 35:172w, 35:182w-35:184w, 35:201, 35:211w, 35:220w, 35:233, 35:251, 35:258w, 35:266, 35:270w, 35:273w, 35:276, 35:285w, 35:292, 35:309w, 35:311w, 35:316, 35:321, 35:323w, 35:335, 35:339, 37:12, 38:01:01, 38:01:03-38:02:05, 38:02:07-38:21, 38:23-38:28, 38:30-38:37, 38:39-38:48, 38:49w, 38:50-38:51, 38:53-38:62, 38:64-38:67, 39:01:01:01-39:01:01:07, 39:01:03-39:01:11, 39:01:12w, 39:01:13-39:01:18, 39:01:20-39:06:05, 39:08-39:14, 39:16, 39:18-39:20, 39:22, 39:24:01-39:24:02, 39:26-39:28, 39:30-39:42, 39:45-39:48, 39:50-39:54, 39:56w, 39:57-39:90, 39:92-39:95N, 39:97N, 39:98w, 39:99-39:108, 39:109w, 39:110-39:128, 41:01:01-41:03:02, 41:05-41:09, 41:11-41:17, 41:19-41:20, 41:22-41:24, 41:25w, 41:26-41:36, 41:38-41:48, 41:49w, 41:50, 42:01:01-42:02:01:02, 42:04-42:08, 42:10-42:15, 42:17-42:19, 42:21-42:24, 44:10, 44:14, 44:62, 44:77, 44:82, 44:166, 44:184, 44:213, 45:14, 46:11, 46:18, 46:61, 48:05, 48:10, 48:12-48:13, 51:01:01:01-51:01:18, 51:01:19w, 51:01:20-51:01:30, 51:01:32-51:01:49, 51:01:51-51:02:02, 51:02:04-51:17, 51:19-51:24:05, 51:26, 51:28-51:36, 51:38-51:41N, 51:43-51:44N, 51:46, 51:48-51:74, 51:76-51:83, 51:86-51:89, 51:91, 51:95-51:121, 51:123-51:129, 51:131-51:134, 51:136-51:145, 51:147-51:162, 51:164-51:171, 51:173Q-51:176, 51:178N-51:194, 51:196, 51:198-51:205, 51:206w, 51:207-51:215, 51:217-51:219, 51:221-51:222, 52:01:01:01-52:01:10, 52:01:12-52:01:19, 52:01:20w, 52:01:21-52:01:22, 52:01:23w, 52:01:24-52:11, 52:13-52:25:02, 52:27-52:56, 52:57w, 52:58-52:62, 52:64-52:69, 53:19w, 53:36w, 54:01:01-54:05N, 54:07-54:08N, 54:10, 54:12-54:13, 54:17-54:21, 54:23-54:26, 54:28-54:38, 55:01:01-55:05, 55:07-55:13, 55:15-55:22, 55:24-55:26, 55:28-55:40, 55:42-55:45, 55:47-55:49, 55:51-55:57, 55:59-55:67, 55:69-55:85, 56:01:01:01-56:02:01, 56:02:02w, 56:04-56:08, 56:10, 56:12-56:17, 56:19N-56:30, 56:33-56:49, 56:51-56:53, 57:02:01-57:03:03, 57:07, 57:09, 57:12, 57:17, 57:28N, 57:39, 57:42w, 57:46, 57:57, 57:63, 57:66, 57:70, 57:80, 57:82, 57:84, 58:08:01-58:08:02, 58:20, 58:27-58:28:02, 58:65w, 58:73, 59:01:01:01-59:09, 67:01:01-67:07, 78:01:01:01-78:09, 82:01-82:03, C*01:90, C*01:116, C*02:68, C*03:317, C*15:24, C*15:128 *07:05:01:01-07:06:02, 07:32, 07:40, 07:53, 07:69, 07:78, 07:90, 07:97, 07:105, 07:112, 07:123, 07:137-07:138, 07:140, 07:176, 07:182N, 07:201N, 07:206-07:207, 07:210, 07:213-07:214, 07:222, 07:249, 07:258, 07:262, 07:264, 07:269-07:270, 07:278, 07:283, 07:287, 07:293, 08:20, 08:40, 08:70, 08:79, 13:18, 13:31, 13:41, 13:73, 15:30, 15:58, 15:73, 15:150, 15:303, 15:324, 35:18, 35:44, 35:172w, 39:33, 39:36, 46:61, 48:05, 48:19, 51:05, 51:29, 51:54, 51:61:01-51:61:02, 51:82, 51:217, 52:21, 52:24, 54:10, 54:20, 54:33, 55:09, 55:21, 55:37, 55:52, 56:43, 56:51, 57:02:01-57:02:02, 57:12, 57:28N, 57:42, 57:63, A*68:02:11, C*01:90, C*01:116, C*02:68 |
| 67 | 85 bp 130 bp 255 bp | *40:89 *40:320 *40:107 | *07:284, 08:74, 41:52 *07:136:01 |
| 68 | 90 bp 130 bp 225 bp | *40:180 *40:320 *40:90 | *08:14, 27:32 *07:136:01 |
| 73 | 220 bp 250 bp | *40:329 *40:176, 40:301 |  |
| 75 | 245 bp 270 bp | *40:171 *40:155:01N-40:155:02N | *15:193, 39:53, C*04:155, C*07:107, C*07:224 |
| 77 | 130 bp 170 bp 195 bp | *40:295 *40:102 *40:142N |  |
| 78 | 115 bp 150 bp 250 bp | *40:111 *40:182, 40:205 *40:135 |  |
| 79 | 180 bp 230 bp | *40:145, 40:268, 40:339 *40:109, 40:116, 40:292 | *18:67 |
| 80 | 115 bp 140 bp | *40:122, 40:151 *40:334 | C*02:98 |
| 81 | 225 bp 325 bp | *40:144N *40:141, 40:299 | *08:163? |
| 82 | 75 bp 120 bp 150 bp 180 bp | *40:115 *40:311 *40:267, 40:346 *40:150, 40:331 | *07:120, 15:175, 15:214, 18:81, 35:250, 51:165, 57:29 |
| 83 | 155 bp 210 bp 240 bp | *40:179, 40:249 *40:104, 40:253, 40:348 *40:247 | *51:72 |
| 84 | 160 bp 205 bp | *40:249 *40:132, 40:178, 40:263N | *51:72 *07:140, 35:309w, 44:10, 44:14, 44:31, 44:62, 44:77, 44:82, 44:184, 44:213, 45:14, 51:23, 51:108, 52:15, 57:07, 82:01-82:03 |
| 85 | 250 bp 355 bp | *40:213, 40:247 *40:337N | *15:227, 27:69, 37:23 *15:26N |
| 87 | 100 bp 195 bp | *40:286N *40:210?, 40:229 | *15:272N *18:43, 27:37 |
| 88 | 105 bp 200 bp | *40:185, 40:209, 40:308, 40:350 *40:201 | *13:58, 18:80, 18:102, 27:52, 35:84, 35:90, 35:150:01-35:150:02, 35:184, 35:201, 44:134, 51:103, 51:119 *13:25, 44:15, 44:109, 44:176 |
| 90 | 220 bp 315 bp 415 bp | *40:329 *40:299 *40:278 | *08:163 *07:282 |
| 91 | 130 bp 335 bp 395 bp | *40:338N *40:303 *40:256N |  |
| 93 | 110 bp 255 bp | *40:310 *40:335 |  |

4The following HLA-B*40 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| B*40:69, 40:221 | 56 |

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Changes in revision R01 compared to R00:

- The reactivities of primer mix 94 have been corrected in the Specificity and Interpretation Tables.

Changes in revision R02 compared to R01:

- Primer mix 46 may give rise to a lower yield of HLA-specific PCR product than the other B*40 primer mixes. A footnote has been added in the Specificity Table.

Changes in revision R03 compared to R02:

- Primer mix 46 may give rise to a most pronounced lower yield of HLA-specific PCR product than the other B*40 primer mixes. The footnote has been changed in the Specificity Table.

Changes in revision R05 compared to R03:

- Primer mix 44 does not amplify the B*40:78 allele. The correction has been implemented in the Specificity and Interpretation tables.