Gel Documentation Form and Worksheet

HLA-B*42 (101.543-06/06u) 	Lot No: 4G2	Expiry Date: 2021-01-01

Name:___________________

Sample ID:_______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:________________

Reviewed By:________________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 16 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-B*42 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.31.0, January 2018.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-B*42 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-B*42 alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- |
| 11 | 115 bp 180 bp | *42:19 *42:10, 42:17 | *08:20, 08:53:01-08:53:02 *08:49, 08:71, 27:83, 41:24, 53:15 |
| 14 | 195 bp 290 bp | *42:15, 42:18 *42:13 | 07:11, 07:57, 07:75:01:01-07:75:01:02, 07:160, 07:240, 07:253, 08:16-08:17, 08:47, 08:123, 08:145, 18:49, 18:79, 18:108, 18:132, 18:142, 37:05, 39:20, 39:124, C*01:59, C*01:118, C*06:82, C*06:210, C*07:49, C*07:124, C*07:155, C*07:210, C*07:238, C*07:247, C*07:403 *08:49, 08:60, 08:76, 08:129, 08:181, 41:48, 53:15 |

4The HLA-B*42 primer set cannot separate the B*42:07 and the B*08:94 alleles. These alleles can be distinguished by the HLA-B low resolution kit and/or the HLA-B*08 high resolution kit.

‘w’, might be weakly amplified.