HLA-B*42 (101.543-06/06u) 	Lot No: 1S2	Expiry Date: 2027-05-01

Name:___________________

Sample ID:_______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:________________

Reviewed By:________________

Interpretation:___________                             Failed lanes: ____________                              Comments:________________

Gel Picture

PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 1, 3 and 12 may give rise to a lower yield of HLA-specific PCR product than the other B*42 primer mixes.

Primer mix 16 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1HLA-B*42 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.34.0, October 2018.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-B*42 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-B*42 alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- |
| 11 | 115 bp 180 bp | *42:19 *42:10, 42:17 | *07:388, 08:20:01-08:20:02, 08:53:01-08:53:02, 08:281, 08:296, 15:516, 15:576 *08:49, 08:71, 08:291, 08:296, 15:516, 15:576, 27:83, 35:530, 41:24, 44:342, 53:15 |
| 14 | 195 bp 290 bp | *42:15, 42:18 *42:13 | *07:11, 07:57, 07:75:01:01-07:75:01:02, 07:160, 07:240, 07:253, 07:340, 07:463, 08:16-08:17, 08:47, 08:123, 08:145, 08:295, 15:530, 18:49, 18:79, 18:108, 18:132, 18:142, 37:05, 39:20, 39:124, 39:168, C*01:59, C*01:118, C*01:157, C*06:82, C*06:210, C*07:49, C*07:124, C*07:155, C*07:210, C*07:238, C*07:247, C*07:403 *08:49, 08:60, 08:76, 08:129, 08:181, 35:530, 41:48, 44:342, 53:15 |
| 15 | 165 bp 230 bp | *42:25 *42:14 | *08:298, 13:61 *08:116, 41:06, 41:15, 51:373, 52:92 |

4The following HLA-B*42 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| B*42:14, 42:25 | 15 |

5The following alleles give rise to identical amplification patterns with the HLA-B*42 high resolution kit. These alleles can be distinguished by the HLA-B low resolution and/or the respective kits.

Alleles

B*42:05:01-42:05:02, 07:349

B*42:07, 42:26, 08:94, 08:226

B*42:19, 07:388

Abbreviations

w: might be weakly amplified.