PAGE  
[image: image2.emf]Well No. 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24

ICB 800 800 800 107010701070107010701070 800 800 10701070 800 1070 800 1070107010701070 800 800 800 1070

260 90 225 160 75 170 155 90 120 125 150 145 90 95 170 145 195 115 150 130 100 205 145 225

255 195 195 150 155 190 180 290

260

Well No. 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24

Specific 

Reaction

AmpS

[image: image3.emf]Well No. 25 26 27 28 29 30 31 32 33 34 35 36 37 38 39 40 41 42 43 44 45 46 47 48

ICB 1070 800 800 1070 800 800 107010701070 800 800 800 1070 800 800 10701070 800 800 800 800 107010701070

145 100 160 130 90 195 110 155 150 160 85 145 205 225 75 235 255 130 110 85 220 165 65 105

135 255 165 255 195 185 185 235 260 165 165 160 250 450 115 140

225 230 140 260

260 195 595

Well No. 25 26 27 28 29 30 31 32 33 34 35 36 37 38 39 40 41 42 43 44 45 46 47 48

Specific 

Reaction

AmpS

[image: image4.emf]Well No. 49 50 51 52 53 54 55 56 57 58 59 60 61 62 63 64

ICB 800 10701070107010701070107010701070107010701070107010701070

85 95 110 130 210 270 255 110 185 210 185 100 255 115 285

160 285 250 300

185 405 280

285

Well No. 49 50 51 52 53 54 55 56 57 58 59 60 61 62 63 64

Specific 

Reaction

NC

AmpS

 ®





HLA-B*44
November 2015
101.563-24/03 -24u/03u

Rev. No: 01                                                                11Y     


Gel Documentation Form and Worksheet 
HLA-B*44 (101.563-24/03, -24u/03u) 
Lot No: 11Y
Expiry Date: 2017-September-01
Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________
 

Interpretation:___________
                       Failed lanes: ___________                          Comments:_________________      
Gel Picture
	PHOTO DOCUMENT


[image: image5.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

HLA-B allele1,2,3

*44:02:01:01, 44:02:01:03-44:02:09, 

44:02:11-44:02:36, 44:113, 44:151-44:152

1 2 3 11 17 23 32

*44:02:01:02S 1 2 3 11 17 2324 32

*44:02:10 1 3 11 17 23 27 32

*44:03:01-44:03:24, 44:109, 44:203 1 2 4 11 17 23

*44:04 1 2 5 11 17 23 35

*44:05:01-44:05:04 1 2 3 6 11 23 32

*44:06 1 3 7 17 ? ? 32

*44:07 1 4 8 11 17 23

*44:08 1 3 8 17 23 32

*44:09 1 2 3 12 17 23 32

*44:10 1 2 6 9 11 ? ? 30 37 40

*44:11 1 2 3 11 17 19 ? ? 32

*44:12 1 2 3 11 14 17 ? ? 32

*44:13 1 2 4 11 13 17 23

*44:14 1 2 3 11 15 ? ? 32 37 42

*44:15 2 6 9 11 25 27 2930 32 34 40 61

*44:16 2 3 11 21 23 3233 39

*44:17 1 2 3 6 11 15 23 ? 3132 37

*44:18 2 6 9 11 16 ? ? 27 30 32 34 40

*44:19N 1 2 3 11 1718 23 32

*44:20, 44:100 2 11 17 ? ? 27 32

*44:21 1 2 3 11 17 21 23 32

*44:22 1 2 3 11 17 20 ? ? 32

*44:23N 1 2 3 9 11 17 23 32

*44:24 1 2 3 1011 17 ? ? 32

*44:25 1 2 3 6 11 16 ? ? 32

*44:26 1 2 4 11 17 2223 ?

*44:27:01-44:27:02 1 2 3 11 17 25 32

*44:28:01-44:28:02 1 2 5 11 17 ? ?

*44:29 1 2 7 11 17 ? ? 43

*44:30 1 2 4 11 17 19 ? ?

*44:31 2 6 11 18 25 37 40

*44:32, 44:1244 1 2 4 11 17 21 ? ?

*44:33 1 2 3 11 17 ? ?26 32

*44:34:01 1 3 11 17 ? ? 27 32

*44:34:02 1 3 11 17 ? 32

*44:35 1 2 11 17 ? ? 28

*44:36 1 2 4 11 17 ? ? 45

*44:37:01-44:37:02 2 4 11 ? ? 33

*44:38 1 2 4 11 17 ? ? 62

*44:39 1 4 6 11 17 23

*44:40, 44:210 1 4 17 ? ? 36

*44:41:01-44:41:02 2 11 17 ? ? 32 38 40

*44:42 2 3 11 ? ? 32 39

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



[image: image6.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*44:88 1 2 3 11 ? ? 32 58

*44:89 1 2 3 7 11 17 ? ? 32 43

*44:90 1 2 3 12 17 ? ? 32 42

*44:91 2 3 11 ? ? 3233 39

*44:92, 44:181 1 2 4 1011 17 ? ?

*44:94, 44:984 1 2 4 11 17 ? ? 31

*44:95 1 2 3 11 1617 ? ? 32

*44:96, 44:153 1 2 4 11 1718 ? ?

*44:97 1 2 3 11 17 ? ? 32 42

*44:99 2 11 17 ? ? 3132

*44:101 2 3 11 23 32

*44:102, 44:112, 44:116, 44:119, 44:121, 

44:162, 44:170, 44:172, 44:176, 44:179, 

44:185, 44:196, 44:206, 44:208, 44:214

1 2 3 11 17 ? ? 32

*44:103, 44:188 1 2 4 11 17 ? ? 29 61

*44:104 1 2 3 11 17 ? ? 32 56

*44:105 1 2 4 11 17 20 ? ?

*44:106 2 11 17 ? ? 32 38 40 45

*44:107 1 2 6 11 ? ? 37

*44:108N 1 2 4 11 17 ? ? 49

*44:110 1 2 4 11 17 23 55

*44:111 1 4 11 17 ? ? 50

*44:114, 44:125, 44:128, 44:141, 44:154-

44:155, 44:157, 44:161, 44:164-44:165, 

44:167, 44:174, 44:178, 44:180, 44:183, 

44:186, 44:189, 44:199, 44:202, 44:207, 

44:209

1 2 4 11 17 ? ?

*44:115 1 2 4 11 17 23 51

*44:117 1 2 3 6 11 23 29 32

*44:118 1 2 3 11 17 23 32 52

*44:120 1 2 4 11 17 ? ? 48

*44:122 1 2 4 11 17 ? ? 53

*44:123 1 2 3 6 11 15 ? ? 32 37

*44:126:01 1 2 3 11 17 ? ? 32 54

*44:126:02 1 2 3 11 17 23 32 54

*44:127, 44:201 1 2 3 17 ? ? 32

*44:130 1 4 17 23 36

*44:131 1 2 3 12 17 ? ? 32

*44:132, 44:190 2 3 11 ? ? 3233 35 39

*44:133 1 2 3 11 ? ? 32

*44:134 1 3 6 ? ? 32

*44:135 2 11 ? ? 37 40 45

*44:137 1 2 3 11 17 23 ? 32 35 50

*44:138Q 2 11 17 25 32 55

*44:140 1 2 3 6 9 11 ? ? 30 32 34

*44:142 1 2 3 11 17 ? ? 32 36

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control

[image: image7.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*44:143 1 2 4 11 17 ? ? 52

*44:144 1 2 6 11 ? ? 31

*44:145, 44:168 1 3 11 17 ? ? 32

*44:146 1 2 11 17 ? ? 44

*44:147, 44:175 1 2 11 17 ? ?

*44:148 1 2 3 11 17 ? ? 32 50 56

*44:149N 1 2 3 11 17 ? ? 32 55

*44:150 2 11 21 23 33 44

*44:156 1 4 17 ? ? 36 41 59

*44:158 2 6 11 ? ? 32 38 40 45

*44:159 1 2 4 11 17 ? ? 54

*44:160Q, 44:198N4 1 2 4 11 17 ? ? 56

*44:163 1 2 4 11 17 ? ? 35

*44:166 2 11 15 ? ? 32 3839 42

*44:171N 1 2 3 11 17 ? ? 32 50

*44:173 1 2 3 11 17 23 32 52 63

*44:177 1 2 3 11 17 2223 ? 32

*44:182 1 2 4 11 1718 ? ? 44

*44:184 2 6 11 23 37 40 45

*44:187 1 2 3 11 17 23 32 60

*44:191 1 2 3 11 17 ? ? 32 4748

*44:193 2 4 11 17 ? ?

*44:194 1 4 8 11 17 ? ? 57

*44:195N 2 3 11 ? ? 32 48 54

*44:200 1 3 11 17 ? ? 32 50

*44:204 1 2 4 11 17 ? ? 47

*44:205 1 2 4 11 17 ? ? 57

*44:211 1 2 3 11 17 23 32 55

*44:213 2 6 11 15 ? ? 32 3738 40 42 45

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control


‘ICB’ Internal Control Band, 
‘AmpS’ Amplicon Size
Notes:
Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 6, 21, 36 and 57 have a tendency to giving rise to primer oligomer formation.

Primer mixes 25, 29, 39, 52 and 55 may have tendencies of unspecific amplifications.

Primer mix 39 may give rise to a lower yield of HLA-specific PCR product than the other B*44 primer mixes.

We assume that the fourth exon nucleotide sequences are conserved within allelic groups. Primer mix 25 amplifies the B*44:27 and B*44:31 alleles. The purpose of primer mixes 23 and 25 is to distinguish e.g. the B*44:02,44:02 and B*44:02,44:27 genotypes.

Primer mix 64 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

[image: image8.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*07:04, 07:19, 07:25, 07:146, 08:07, 

08:37, 08:77, 08:89, 08:107-08:108, 

15:51, 15:179:01-15:179:02, 15:199, 

15:218Q, 37:09, 42:09, 42:16

15 25 32 3839

*07:20, 07:140, 07:153 5 6 25 37 40 45

*07:28 21 25 33 44

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



25



*07:02:01-07:03, 07:05:01-07:18:02, 

07:21-07:24, 07:26-07:27, 07:29-07:57, 

07:59-07:132, 07:134-07:139, 07:141-

07:145, 07:147-07:148, 07:150-07:152, 

07:154-07:241, 08:13, 08:20, 08:25, 

08:40, 08:55, 08:70, 08:79, 13:09, 

13:12:02, 13:30, 13:39, 13:67, 

15:01:01:01-15:01:04, 15:01:06-15:13:02, 

15:15-15:19, 15:21, 15:23-15:32:01, 15:33-

15:35, 15:37-15:40, 15:42-15:45, 15:47:01-

15:50, 15:52, 15:54-15:58, 15:60-15:76, 

15:78:01-15:82, 15:84-15:88, 15:90, 15:92-

15:99, 15:101, 15:103-15:105, 15:107-

15:114, 15:116-15:129, 15:132-15:142, 

15:144-15:160, 15:162-15:178, 15:180-

15:182N, 15:184-15:198, 15:200, 15:202-

15:211, 15:213-15:217, 15:219-15:227, 

15:229-15:232, 15:234-15:246N, 15:248-

15:255, 15:257-15:270, 15:272N-15:283, 

15:285-15:304N, 15:306-15:338, 15:340-

15:343, 18:01:01:01-18:08, 18:10-18:15, 

18:17N-18:22, 18:24-18:47, 18:49-18:55, 

18:57-18:66, 18:68-18:79, 18:81-18:101, 

18:103-18:106, 27:02:01-27:21, 27:23-

27:46, 27:48-27:49, 27:51, 27:53-27:66N, 

27:69-27:126, 27:128-27:135, 35:31, 

35:43:01, 35:46, 35:57, 35:60, 35:79, 

35:185, 35:279, 37:02, 37:07, 37:39-

37:40, 39:06:05, 39:09:03, 39:100, 40:12, 

40:21, 40:46, 40:93, 40:137, 40:265N, 

42:19, 46:01:01-46:09, 46:11-46:16, 46:18-

46:63, 48:01:01-48:01:06, 48:03:01-48:36, 

49:26, 54:01:01-54:34, 55:01:01-55:01:02, 

55:01:04-55:05, 55:07-55:19, 55:21-

55:55N, 55:57-55:70, 56:01:01:01-56:04, 

56:06, 56:08-56:12, 56:14-56:20:02, 56:22-

56:45, 57:01:01-57:03:02, 57:05-57:06, 

57:08, 57:10-57:23, 57:25-57:44, 57:46-

57:50, 57:52-57:68, 57:70-57:74, 

59:01:01:01-59:07, 81:01-81:02, 81:04N-

81:07

 

[image: image1.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*44:43:01-44:43:02 1 2 6 11 ? ? 31 37

*44:44 1 3 17 ? ? 32 36

*44:45 1 2 11 ? ? 37

*44:46, 44:129 1 2 4 12 17 23

*44:47, 44:197 2 4 11 17 ? ? 27

*44:48 1 2 3 11 17 ? ? 32 40

*44:49 1 2 3 11 17 23 32 41

*44:50:01 1 2 4 11 1617 ? ?

*44:50:02 1 2 4 11 1617 23

*44:51 1 2 3 11 15 23 32 37

*44:52N 1 2 3 11 17 23 32 43

*44:53 1 2 3 11 17 23 32 44

*44:54 2 4 11 17 ? ? 40 45

*44:55, 44:212 1 2 3 11 17 ? ? 29 32 61

*44:56N 1 2 5 11 17 23 35 46

*44:57 1 4 8 17 23

*44:58N, 44:1394 1 2 3 6 11 ? ? 32 47

*44:59:01 1 2 3 11 17 23 32 48

*44:59:02 1 2 3 11 17 ? 32 48

*44:60 1 4 17 ? ? 57

*44:61N 1 2 4 11 17 23 48

*44:62 1 2 3 6 11 15 ? ? 32 37 42

*44:63 1 2 3 11 17 ? ? 32 35

*44:64:01 2 4 11 21 23 ? 33

*44:64:02 2 4 11 21 23 33

*44:65 2 4 11 17 23 40

*44:66 1 2 3 11 17 23 32 46

*44:67 1 2 3 11 13 17 ? ? 32

*44:68 1 2 3 11 17 ? ? 28 32

*44:69, 44:854 1 2 4 11 17 ? ?26

*44:70 1 2 3 5 6 11 ? ? 32

*44:71 1 2 3 11 17 ? ? 32 62

*44:72, 44:934 1 2 3 11 17 ? ? 32 47

*44:73 1 2 3 11 17 ? ? 32 44

*44:74, 44:169 1 2 3 11 1718 ? ? 32

*44:75 1 2 3 6 12 ? ? 32

*44:76 1 2 5 11 17 ? ? 44

*44:77 1 2 6 11 23 37

*44:78, 44:1364 1 2 3 6 11 ? ? 32 48

*44:79:01-44:79:02, 44:192 1 2 4 11 17 ? ? 44

*44:80 1 2 3 11 17 ? ? 3233

*44:81 1 4 11 17 23 33

*44:82 1 2 6 11 18 ? ? 37

*44:83 1 3 17 23 32 57

*44:84:01 1 2 3 11 17 ? ? 32 57

*44:84:02 1 2 3 11 17 ? ? 32 59

*44:86 1 2 3 11 17 ? ? 32 45

*44:87 1 2 3 11 17 ? ? 32 58

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



[image: image9.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*07:133 12 w 25

*07:149 25 36 57

*08:01:01-08:01:22, 08:01:24-08:01:26, 

08:01:28-08:01:31, 08:03-08:05, 08:08N, 

08:10-08:12:03, 08:15-08:19N, 08:21-

08:24, 08:26:01-08:26:02, 08:29-08:36, 

08:38-08:39, 08:41-08:42, 08:44-08:48, 

08:50-08:54, 08:56:01-08:59:02, 08:61-

08:69, 08:71-08:73, 08:75, 08:78, 08:80-

08:83, 08:85, 08:88, 08:90-08:103, 08:105-

08:106, 08:109-08:110, 08:112-08:116, 

08:118-08:128, 08:130-08:134, 08:136-

08:140, 35:87, 42:01:01-42:02, 42:05:01-

42:08, 42:10-42:12, 42:14-42:15, 42:17-

42:18

15 25 32 3839 42

*08:01:23, 08:14, 08:28, 08:74, 08:111, 

37:44, 37:51, 42:20

25 32 3839

*08:01:27, 08:86N, 08:135 15 25 32 38 42

*08:02, 08:117 15 25 32 36 3839 4142 59

*08:09, 08:27, 08:84, 08:104, 15:83, 

42:04, 55:20

15 25 32 39 42

*08:43 15 25 38 42

*08:49, 37:01:01-37:01:07, 37:01:09, 

37:03N-37:05, 37:08, 37:11, 37:13-37:15, 

37:17-37:18, 37:20-37:31, 37:33N, 37:36-

37:38, 37:41-37:43, 37:45-37:50, 37:52, 

37:54

3 25 32 39

*08:60, 08:76, 08:129, 37:12, 42:13, 57:09

3 15 25 32 39 42

*08:87 25 3233 35 3839 44

*13:01:01-13:04, 13:06-13:08, 13:10-

13:12:01, 13:13-13:22:02, 13:25-13:29, 

13:31-13:38, 13:40-13:45, 13:47-13:54, 

13:56:01N-13:57, 13:59-13:66, 13:68-

13:79, 13:82-13:83, 15:36, 15:89, 15:256, 

15:339

25 57

*13:23, 13:55 25 29 57 61

*13:46 25 32 39 57

*13:58 8 25 57

*13:80 25 57 62

*13:81 22 25 57

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



[image: image10.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*14:01:01-14:09, 14:11-14:43, 15:20, 

35:01:01:01-35:03:18, 35:04:01-35:15:01, 

35:16-35:18, 35:20:01-35:30, 35:32:01-

35:37, 35:39-35:42:02, 35:44, 35:48-

35:56, 35:58, 35:61:01-35:62, 35:64:01-

35:70, 35:72, 35:74-35:78, 35:80-35:83, 

35:85-35:86, 35:88-35:89, 35:91-35:114, 

35:116-35:137, 35:139-35:149, 35:151-

35:153, 35:155-35:167, 35:170-35:182, 

35:186-35:200, 35:202-35:203, 35:205-

35:216N, 35:218-35:222, 35:224-35:229, 

35:231-35:241, 35:243-35:246, 35:248-

35:259, 35:261-35:273, 35:275-35:278, 

38:01:01-38:01:07, 38:02:07, 38:05, 38:09-

38:14, 38:16-38:17, 38:19, 38:21-38:22, 

38:24-38:28, 38:31-38:34N, 38:36-38:42, 

38:52-38:56, 39:01:01:01-39:01:01:03, 

39:01:03-39:06:04, 39:07-39:09:02, 

39:10:01-39:20, 39:22-39:41, 39:43-

39:45, 39:47-39:68, 39:70-39:99, 48:02:01-

48:02:02, 51:01:40, 51:104, 51:147, 

52:02:02, 52:16, 52:32, 53:03, 56:05:01-

56:05:02, 56:21, 58:17N, 58:36, 58:44, 

58:64, 67:01:01-67:04, 78:01:01-78:08, 

81:03

w

*14:10, 35:59, 35:242, 38:20, 39:42 w 29

*15:14, 15:91, 15:131, 15:161, 46:17 6 25 31 37 40

*15:32:02 25 27

*15:46, 15:106, 40:10:01-40:10:02, 

40:149, 40:177, 40:215, 40:248, 40:282, 

40:294

12 25

*15:53, 15:212, 40:01:01-40:01:10, 

40:01:12-40:07, 40:09, 40:11:01-40:11:02, 

40:14:01-40:16, 40:18, 40:20, 40:22N-

40:24, 40:26-40:40, 40:42-40:45, 40:48-

40:57, 40:59-40:67, 40:69-40:75, 40:77-

40:82, 40:84-40:92, 40:94-40:95, 40:97-

40:105, 40:107-40:108, 40:111-40:116, 

40:118N-40:131, 40:133Q-40:135, 40:138-

40:148, 40:150-40:156, 40:158-40:165, 

40:167-40:169, 40:171-40:176, 40:179-

40:184, 40:186-40:187, 40:189-40:200, 

40:202-40:207, 40:210-40:214, 40:216N-

40:230, 40:233-40:247, 40:249-40:264, 

40:266-40:278, 40:280-40:281, 40:283-

40:291N, 40:293, 40:295-40:298, 47:02, 

50:01:01-50:01:06, 50:04-50:05, 50:08-

50:14, 50:16-50:20, 50:31-50:36, 50:38-

50:42

2 12 25

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



[image: image11.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*38:06-38:07, 51:01:01:01, 51:01:02-

51:01:39, 51:01:41-51:07:02, 51:09:01-

51:19, 51:21-51:22, 51:24:01-51:24:05, 

51:26-51:35, 51:37-51:41N, 51:43, 51:45-

51:46, 51:48-51:50, 51:52-51:53, 51:55-

51:77, 51:79-51:96, 51:98N-51:102, 

51:105-51:107, 51:109-51:111, 51:113-

51:118N, 51:120-51:140, 51:142-51:146, 

51:148-51:152, 51:154-51:156, 51:158:01-

51:184N, 52:01:01:01-52:01:01:03, 

52:01:03, 52:01:05-52:01:08, 52:01:10-

52:01:21, 52:04-52:05, 52:08, 52:10:01-

52:10:02, 52:11-52:14, 52:17-52:18, 52:22-

52:25, 52:27-52:31:01, 52:33-52:44, 

53:01:01-53:02, 53:04-53:08:02, 53:10, 

53:14-53:21, 53:23-53:26, 53:28-53:29, 

53:32-53:35, 53:37, 58:01:01-58:02, 58:04-

58:06, 58:08:01-58:10N, 58:12-58:16:02, 

58:18-58:19, 58:21-58:29, 58:31N-58:35, 

58:37-58:43, 58:46-58:63, 58:65-58:69

7 w

*38:18 36 59

*38:30 15 w 32 3839 42

*38:51 w26

*39:46 w 52

*40:01:11, 40:58 12 25 27

*40:08, 40:25, 40:68, 40:106, 40:232, 

50:15

2 12 14 25

*40:13, 40:19, 40:109, 40:117, 40:292, 

49:01:01-49:01:07, 49:04:01-49:14, 49:16-

49:19N, 49:22-49:23, 49:28-49:35, 57:51, 

57:69

2 11 16 25

*40:47, 40:96, 40:157, 40:201, 49:02 2 11 25 29 61

*40:76, 40:208, 50:06 2 25

*40:83, 40:170, 50:07 8 12 25

*40:110, 40:188, 47:01:01:01-47:01:02, 

47:04-47:06, 47:08-47:09

2 11 25

*40:132, 40:178 2 6 12 18 25 37 40

*40:136, 40:231, 40:279, 41:02:01-

41:03:02, 41:08, 41:11, 41:13, 41:15, 

41:19, 41:23-41:24, 41:27, 41:31

2 12 15 25 32 3839 42

*40:166 2 12 w 25 42

*41:01:01-41:01:03, 41:06-41:07, 41:09, 

41:14, 41:16-41:17, 41:20, 41:25, 41:28-

41:30, 41:32-41:35

2 12 15 25 32 39 42

*41:04, 41:18 2 12 15 25 32 3839

*41:05, 41:26 2 15 25 32 39 42

*41:10 2 12 25 32 3839

*41:12, 41:22 12 15 25 32 39 42

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



[image: image12.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*41:21 2 12 15 25 32 39

*45:01:01-45:01:02, 45:05, 45:11-45:13, 

45:15

2 6 9 12 25 27 30 32 34 40

*45:02 2 6 12 25 27 32 40

*45:03 2 12 25 27 32 40

*45:04, 45:08, 45:10 2 12 25 32 34 39

*45:06 2 6 9 12 25 27 30 32 34 40 42

*45:07 2 6 9 12 w 27 30 32 34 40

*45:09 2 6 12 15 25 27 3132 37 40

*45:14 2 6 9 12 15 25 27 30 32 34 37 40 42

*47:03 2 11w 25

*47:07 2 11 22 25

*49:03 7 25 27

*49:20 2 6 9 11 16 25 27 30 40

*49:21 2 16 25

*49:24 2 11 16 20 25

*49:25, 49:27, 57:45 11 16 25

*50:02 2 6 9 12 25 27 30 40

*50:37 2 12

*51:01:01:02, 51:51, 51:185 7

*51:08:01-51:08:02, 51:20, 51:36, 51:44N, 

51:97, 51:141, 51:153, 52:19

7 15 w 32 39 42

*51:23, 51:108, 52:15 6 7 w 27 37

*51:42 1 4 7 17 23

*51:54, 51:78:01-51:78:02, 52:20, 53:09, 

53:11-53:13, 53:31, 53:36, 53:38

w 57

*51:103, 51:119 7 8 w

*51:112, 53:30, 53:39 11 14 16 w

*52:01:02, 52:01:04, 52:01:09, 52:03, 

52:06:01-52:06:02, 52:09, 52:10:03, 

52:21, 52:31:02

7 w 27

*52:02:01 w 27

*52:26 4 7 w 33 44

*53:22 1 3 6 7 15 w 3132 37 40

*53:27 7 w 45

*55:56, 56:13 15 25 32 34 39 42

*57:04 25 33

*57:07 1 6 18 25 37

*58:07 6 7 w 31 37 40

*58:45:01-58:45:02 7 w 54

*82:01-82:02:02 6 9 15 25 30 32 34 3738 40 42

*82:03 6 9 15 ? 30 32 34 3738 40 42

A*02:432, A*30:29, A*31:61 31

A*23:31 4 33

A*24:106 33

A*24:137 26

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control



[image: image13.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

C*02:82 6 31 40

C*03:96, C*15:52 54

C*03:147 48

C*03:251, C*05:10, C*08:44,C*08:61, 

C*08:82

50 56

C*03:265N, C*15:96Q 56

C*05:62, C*07:68, C*07:140, C*07:151, 

C*07:199:01-07:199:02, C*07:364

3839

C*05:102 5 6 40

C*07:04:01-07:04:10, C*07:11-07:12, 

C*07:45, C*07:63, C*07:101, C*07:139, 

C*07:142, C*07:181, C*07:272, C*07:302, 

C*07:323-07:324, C*07:328-07:329N, 

C*07:336, C*07:338, C*07:354-07:355, 

C*07:357, C*07:361, C*07:365, C*07:378, 

C*07:394-07:395, C*07:403, C*07:406

39

C*07:69, C*08:68 42

C*07:231 4 33 35 44

C*08:93 6 40

C*14:61 5 6 31 40

C*15:05:08, C*17:01:10 29

C*15:25 1 3 6 40

C*16:10 33 35 44

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control


[image: image14.emf]Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

*15:77, 15:233 25 33 44

*15:115, 56:07, 59:08 25 36 59

*15:143, 18:48 12 14 25

*15:183, 18:80, 18:102, 27:52, 27:67, 

40:185, 40:209, 46:10

8 25

*15:247 23

*18:09 25 29 57 61

*18:23N 25 49

*18:56 6 15 25 3132 3738 40

*18:67, 51:157, 52:07, 58:11, 58:20 7 25

*27:01 25 29 36 61

*27:47 23 32 38 40

*27:50:01-27:50:02 25 40

*27:127, 83:01 1 3 17 25 32

*35:19, 35:47, 35:154, 35:274, 39:69 12 14 w

*35:38, 35:115, 35:260 3 15 w 32 39

*35:45 1 5 6 w 31 37 40

*35:63 2 12 w

*35:71 1 6 w 31 37 40

*35:84, 35:90, 35:150:01-35:150:02, 

35:184, 35:201

8 w

*35:138, 35:183 22 w

*35:169 1 3 6 15 w 32 37 40

*35:204, 35:223 w 48

*35:217 2 12 14 w

*35:247 w 54

*37:01:08 25 32 39

*37:06:01-37:06:02 3 11 25 32 39

*37:10 3 25 29 32 39 57 61

*37:16Q 3 25 32 39 56

*37:19 3 25 29 32 39

*37:32, 37:53 3 25 32

*37:34 3 7 25 32 39

*37:35, 57:24 3 15 25 32 39

*38:02:01-38:02:06, 38:03-38:04, 38:08, 

38:15, 38:23, 38:29, 38:35, 38:43-38:50

w 36 59

Well No. 1 2 3 4 5 6 7 8 910111213141516171819202122232425262728293031323334353637383940414243444546474849505152535455565758596061626364

Negative Control


[image: image15.png]







1HLA-B*44 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.19.0, January 2015.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 5: Specific PCR fragment of 75 bp in the B*44:04, 44:28:01-44:28:02, 44:56N and 44:76 and the B*07:20, 07:140, 07:153, 35:45 and in the C*05:102 and C*14:61alleles. Specific PCR fragment of 255 bp in the B*44:70 allele.
Primer mix 6: Specific PCR fragment of 170 bp in the B*44:39 allele. Specific PCR fragment of 195 bp in the B*44:05:01-44:05:04, 44:10, 44:15, 44:17-44:18, 44:25, 44:31, 44:43:01-44:43:02, 44:58N, 44:62, 44:70, 44:75, 44:77-44:78, 44:82, 44:107, 44:117, 44:123, 44:134, 44:136, 44:139-44:140, 44:144, 44:158, 44:184 and 44:213 and the B*07:20, 07:140, 07:153, 15:14, 15:91, 15:131, 15:161, 18:56, 35:45, 35:71, 35:169, 40:132, 40:178, 45:01:01-45:02, 45:05-45:07, 45:09, 45:11-45:15, 46:17, 49:20, 50:02, 51:23, 51:108, 52:15, 53:22, 57:07, 58:07 and 82:01-82:03 and in the C*02:82, C*05:102, C*08:93, C*14:61 and C*15:25 alleles.

Primer mix 9: Specific PCR fragment of 120 bp in the B*44:23N allele. Specific PCR fragment of 195 bp in the B*44:10, 44:15, 44:18 and 44:140 and the B*45:01:01-45:01:02, 45:05-45:07, 45:11-45:15, 49:20, 50:02 and 82:01-82:03 alleles.

Primer mix 10: Specific PCR fragment of 125 bp in the B*44:24 and 44:181 alleles. Specific PCR fragment of 150 bp in the B*44:92 allele. 

Primer mix 18: Specific PCR fragment of 115 bp in the B*44:31, 44:169 and 44:182 and in the B*40:132 and 40:178 alleles. Specific PCR fragment of 155 bp in the B*44:74 allele. Specific PCR fragment of 260 bp in the B*44:19N, 44:82, 44:96 and 44:153 and in the B*57:07 alleles.

Primer mix 19: Specific PCR fragment of 150 bp in the B*44:11 allele. Specific PCR fragment of 190 bp in the B*44:30 allele.

Primer mix 21: Specific PCR fragment of 100 bp in the B*44:32 allele. Specific PCR fragment of 180 bp in B*44:16, 44:21, 44:64:01-44:64:02, 44:124 and 44:150 and the B*07:28 alleles.
Primer mix 22: Specific PCR fragment of 205 bp in the B*44:177 and the B*13:81 and 35:183 alleles. Specific PCR fragment of 290 bp in the B*44:26 and the B*35:138 and 47:07 alleles.  

Primer mix 26: Specific PCR fragment of 100 bp in the B*44:69 and the B*38:51 alleles. Specific PCR fragment of 135 bp in the B*44:33 allele. Specific PCR fragment of 225 bp in the B*44:85 and the A*24:137 alleles.

Primer mix 27: Specific PCR fragment of 160 bp in the B*44:02:10 and 44:34:01 and the B*15:32:02, 40:01:11, 40:58, 49:03, 52:01:02, 52:01:04, 52:01:09, 52:02:01, 52:03, 52:06:01-52:06:02, 52:09, 52:10:03, 52:21 and 52:31:02 alleles. Specific PCR fragment of 255 bp in the B*44:15, 44:18, 44:20, 44:47, 44:100 and 44:197 and the B*45:01:01-45:03, 45:05-45:07, 45:09, 45:11-45:15, 49:20, 50:02, 51:23, 51:108 and 52:15 alleles.

Primer mix 28: Specific PCR fragment of 130 bp in the B*44:68 allele. Specific PCR fragment of 165 bp in the B*44:35 allele.

Primer mix 29: Specific PCR fragment of 90 bp in the B*44:117 and the B*14:10, 35:59, 35:242, 37:19, 38:20 and 39:42 and in the C*15:05:08 and C*17:01:10 alleles. Specific PCR fragment of 255 bp in the B*44:15, 44:55, 44:103, 44:188 and 44:212 and the B*13:23, 13:55, 18:09, 27:01, 37:10, 40:47, 40:96, 40:157, 40:201 and 49:02 alleles.

Primer mix 31: Specific PCR fragment of 110 bp in the B*44:94 allele. Specific PCR fragment of 195 bp in the B*44:17, 44:43:01-44:43:02 and 44:144 and the B*15:14, 15:91, 15:131, 15:161, 18:56, 35:45, 35:71, 45:09, 46:17, 53:22 and 58:07 and in the C*02:82 and C*14:61 alleles. Specific PCR fragment of 230 bp in the B*44:98 and in the A*02:432, A*30:29 and A*31:61 alleles. Specific PCR fragment of 260 bp in the B*44:99 allele.

Primer mix 33: Specific PCR fragment of 150 bp in the B*44:81 allele. Specific PCR fragment of 185 bp in the B*44:16, 44:37:01-44:37:02, 44:64:01-44:64:02, 44:80, 44:91, 44:132, 44:150 and 44:190 and the B*07:28, 08:87, 15:77, 15:233, 52:26 and 57:04 and in the A*23:31, A*24:106, C*07:231 and C*16:10 alleles.

Primer mix 35: Specific PCR fragment of 85 bp in the B*44:63 allele. Specific PCR fragment of 185 bp in the B*44:04, 44:56N, 44:132, 44:137, 44:163 and 44:190 and the B*08:87 and in the C*07:231 and C*16:10 alleles.
Primer mix 36: Specific PCR fragment of 145 bp in the B*44:40, 44:44, 44:130, 44:156 and 44:210 and the B*07:149, 08:02, 08:117, 15:115, 27:01, 38:02:01-38:02:06, 38:03-38:04, 38:08, 38:15, 38:18, 38:23, 38:29, 38:35, 38:43-38:50, 56:07 and 59:08 alleles. Specific PCR fragment of 235 bp in the B*44:142 allele.

Primer mix 40: Specific PCR fragment of 235 bp in the B*44:10, 44:15, 44:18 and 44:48 and the B*35:45, 35:71, 35:169, 45:01:01-45:03, 45:05-45:07, 45:09, 45:11-45:15, 49:20, 50:02, 53:22 and in the C*15:25 alleles. Specific PCR fragment of 260 bp in the B*44:31, 44:41:01-44:41:02, 44:54, 44:65, 44:106, 44:135, 44:158, 44:184 and 44:213 and the B*07:20, 07:140, 07:153, 15:14, 15:91, 15:131, 15:161, 18:56, 27:47, 27:50:01-27:50:02, 40:132, 40:178 and 46:17 alleles. Specific PCR fragment of 235 and 260 bp in the B*58:07 and 82:01-82:03 and in C*02:82, C*05:102, C*08:93 and C*14:61 alleles.

Primer mix 42: Specific PCR fragment of 130 bp in the B*44:90 and 44:97 and the B*40:166 and 45:06 and the C*07:69 and C*08:68 alleles. Specific PCR fragment of 165 bp in the B*44:14, 44:62, 44:166 and 44:213 and the B*08:01:01-08:01:22, 08:01:24-08:05, 08:08N-08:12:03, 08:15-08:19N, 08:21-08:24, 08:26:01-08:27, 08:29-08:36, 08:38-08:39, 08:41-08:48, 08:50-08:54, 08:56:01-08:69, 08:71-08:73, 08:75-08:76, 08:78, 08:80-08:86N, 08:88, 08:90-08:106, 08:109-08:110, 08:112-08:140, 15:83, 35:87, 37:12, 38:30, 40:136, 40:231, 40:279, 41:01:01-41:03:02, 41:05-41:09, 41:11-41:17, 41:19-41:20, 41:22-41:35, 42:01:01-42:02, 42:04-42:08, 42:10-42:15, 42:17-42:18, 45:14, 51:08:01-51:08:02, 51:20, 51:36, 51:44N, 51:97, 51:141, 51:153, 52:19, 55:20, 55:56, 56:13, 57:09 and 82:01-82:03 alleles.
Primer mix 43: Specific PCR fragment of 110 bp in the B*44:52N allele. Specific PCR fragment of 165 bp in the B*44:29 and 44:89 alleles.

Primer mix 44: Specific PCR fragment of 85 bp in the B*44:53, 44:182 and 44:192 and the B*08:87 and in the C*07:231 and C*16:10 alleles. Specific PCR fragment of 160 bp in the B*44:73, 44:76, 44:79:01-44:79:02, 44:146 and 44:150 and the B*15:77, 15:233 and 52:26 alleles. Specific PCR fragment of 85 bp and 160 bp in the B*07:28 allele.

Primer mix 45: Specific PCR fragment of 220 bp in the B*44:36 allele. Specific PCR fragment of 250 bp in the B*44:54, 44:86, 44:106, 44:135, 44:158 and 44:184 and the B*07:20, 07:140, 07:153 and 53:27 alleles.

Primer mix 46: Specific PCR fragment of 165 bp in the B*44:66 allele. Specific PCR fragment of 450 bp in the B*44:56N allele.

Primer mix 47: Specific PCR fragment of 65 bp in the B*44:93 and 44:139 alleles. Specific PCR fragment of 115 bp in the B*44:72 and 44:204 alleles. 
Specific PCR fragment of 140 bp in the B*44:191 allele. Specific PCR fragment of 195 bp in the B*44:58N allele.

Primer mix 48: Specific PCR fragment of 105 bp in the B*44:61N allele. Specific PCR fragment of 140 bp in the B*44:191 and the B*35:223 alleles. Specific PCR fragment of 260 bp in the B*44:78, 44:120 and 44:195N and the B*35:204 and in the C*03:147 alleles. Specific PCR fragment of 140 bp and 595 bp in the B*44:59:01-44:59:02 and 44:136 alleles. 
Primer mix 50: Specific PCR fragment of 95 bp in the B*44:111 and 44:200 alleles. Specific PCR fragment of 160 bp in the B*44:171N allele. Specific PCR fragment of 185 bp in the B*44:137 allele. Specific PCR fragment of 285 bp in the B*44:148 and the C*03:251 and in the C*05:10, C*08:44, C*08:61 and C*08:82 alleles. 
Primer mix 52: Specific PCR fragment of 130 bp in the B*44:143 allele. Specific PCR fragment of 285 bp in the B*44:173 allele.  Specific PCR fragment of 405 bp in the B*44:118 and the B*39:46 alleles.
Primer mix 56: Specific PCR fragment of 110 bp in the B*44:198N and in the C*03:265N alleles. Specific PCR fragment of 250 bp in the B*44:160Q and the B*37:16Q and in the C*15:96Q alleles. Specific PCR fragment of 280 bp in the B*44:104 and 44:148 and in the C*03:251, 05:10, C*08:44, C*08:61 and C*08:82 alleles. 

Primer mix 62: Specific PCR fragment of 115 bp in the B*44:38 and the C*13:80 alleles. Specific PCR fragment of 300 bp in the B*44:71 allele.  
4The following HLA-B*44 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

	Alleles
	Primer mix
	Alleles
	Primer mix

	B*44:32, 44:124
	21
	B*44:78, 44:136
	48

	B*44:58N, 44:139
	47
	B*44:94, 44:98
	31

	B*44:69, 44:85
	26
	B*44:160Q, 44:198N
	56

	B*44:72, 44:93
	47


The HLA-B*44 subtyping cannot distinguish the following silent mutations: the B*44:02:01:01, 44:02:01:03-44:02:09 and 44:02:11-44:02:36 alleles, the B*44:03:01-44:03:24 alleles, the B*44:05:01-44:05:04 alleles, the B*44:27:01-44:27:02 allele,  the B*44:28:01-B*44:28:02 alleles, the B*44:37:01-44:37:02 allele,  the B*44:41:01-44:41:02 allele, the B*44:43:01-44:43:02 alleles or the B*44:79:01-44:79:02 alleles. 

‘w’, might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.

Change in revision R01 compared to R00:

1. Primer mix 57 amplifies the B*18:09 allele. This has been corrected in the Specificity and Interpretation Tables. 
Olerup SSP AB

For In Vitro Diagnostic Use 
Franzengatan 5


SE-112 51 Stockholm

Sweden


Page 1 of 15

