PAGE  
Page 1(3)

December 2007
HLA-B*46
Rev. No: 00
Y24



	PHOTO DOCUMENT


Olerup SSP AB
Olerup SSPTM HLA-B*46
Prod. No:
101.544-06 – including Taq polymerase

101.544-06u – without Taq polymerase
Lot No:    
Y24
Expiry Date:    
2009-September-01
Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ug/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 14 primer mixes used for HLA-B*46 SSP typing.
	Primer Mix
	Size of spec. PCR product1
	Size of control band2
	Amplified HLA-B*46 alleles
	Other amplified HLA Class I alleles3

	  14
	130 bp
	  800 bp
	460101-4603, 4605-4612
	1557weakly

	  24
	125 bp
	1070 bp
	4602
	

	  34
	115 bp
	1070 bp
	4603
	1318, 1403, 1547, 1549, 1565, 1584, 1599, 1813, 350801-350803, 3545, 3561, 3566, 3908weakly, 3911weakly, 3933, 4404, 4428, 4456N, 5105, 5129, 5311, 5410, 5509, 5702, 5704, 5705, 5712

	  44
	130 bp
	1070 bp
	4604
	

	  5
	235 bp
	1070 bp
	460101-4604, 4606-4612
	0815, 1557weakly, 3574, 4073, 5503

	  6
	230 bp
	  800 bp
	4606
	0709, 0711, 9538, 3566, 4038, 4052, 4059, 4209, 4814

	  74
	140 bp
	1070 bp
	4607N
	1407N, 3940N, 5619N

	  84,5
	115 bp
	1070 bp
	4608, 4611
	1301-1304, 1306, 1307N, 1309-1317, 1319, 1320, 1401-140202, 1404-1407N, 1502, 1503, 1505, 1506, 1509, 1510, 1513, 1516-1518, 1520, 1521, 1523, 1525, 1529, 1531, 1536, 1537, 1539, 1540, 1542, 1544, 1548, 1552, 1555, 1561, 1562, 1564, 1567, 1569, 1574, 1580, 1586, 1588, 1589, 1591, 1593, 1595, 1598, 9503, 9506, 9508, 9512, 9514, 9515, 9519, 9521, 9523, 9524, 9527, 9531-9534, 9536, 9538, 9539, 180101-1812, 1814, 1815, 1817N, 1818, 1820, 1822-1827, 2701-270402, 270502-2706, 2708-2710, 2712-2719, 2721, 2723, 2726-2732, 2735, 2737, 2738,  350101, 350103-3507, 3510-3513, 3515-3517, 3519-3530, 3532-3537, 3539-3542, 3546-3560, 3563, 3565Q, 3568-3572, 3574-3578, 3582, 3583, 3702, 380101-3817, 39010101-39010102L, 390103-3920, 3922-3932, 3934, 3935, 3937-3943, 4026, 4028, 4035, 4071, 4077, 440301-440303, 4407, 4410, 4413, 4426, 4429, 4430, 4432, 4436-4440, 4443, 4445-4447, 4450, 4454, 4457, 47010101-4705, 4802, 4812, 4814, 4901-4905, 5001, 5002, 5004, 510101-5104, 5106, 5107, 510901-510902, 5111N-5119, 5121-5123, 5126-5128, 5130-5135, 5137-5143, 5145, 5146, 5148-5151, 520101-5211, 530101-5310, 5312, 5313, 5401-5405N, 5407-5409, 5412, 5413, 550101-5503, 5505, 5507, 5510-5519, 5522-5528, 5601, 5602, 5604-5612, 5614-5619N, 5621, 570101-570103, 570301-570302, 5706-5708, 5710, 5711, 5714,  5801, 5802, 5804-5815, 5901, 670101-6702, 7801-7805,

A*3304

	  94
	115 bp
	1070 bp
	460101-4602, 4604, 4605, 4607N, 4609, 4610, 4612
	15010101-150104, 150106, 1504, 1507, 1508, 151101-1512, 1514, 1515, 1519, 1524, 1526N-1528, 1530, 1532, 1534, 1535, 1538, 1543, 1545, 1546, 1550, 1553, 1554, 1556-1558, 1560, 1563, 1566, 1568, 1570, 1571, 1573, 1575-1577, 1579N, 1581, 1582, 1585, 1587, 1592, 1594N, 1596, 1597, 9501, 9502, 9504, 9505, 9509-9511N, 9513, 9517, 9518, 9520, 9522, 9525, 9526, 9528, 9529, 9535, 9537, 1819, 2725, 351401-351402, 3543, 3544, 3562, 3567, 3579 weakly, 3918, 3936, 5406, 5521, 5603

	104
	100 bp
	  800 bp
	4605
	9525, 3554

	11
	150 bp
	1070 bp
	4609
	3562

	12
	170 bp
	1070 bp
	4610
	

	134
	105 bp
	1070 bp
	4611
	130201-1303, 1308Q, 1309, 1314-1316, 1318, 1319, 1542, 3560, 4415, 4418, 4501, 4503-4507, 4901-4903, 5001, 5002, 5004, 5115, 5401-5403, 5405N, 5407, 5408N, 5410, 5412, 5413, 550101-5503, 5505, 5507, 5509-5512, 5515, 5516, 5518, 5519, 5521, 5522, 5524-5526, 5601, 5607, 5608, 5613, 5614, 5616, 5617, 5619N, 5620, 5901 

	14
	215 bp
	1070 bp
	4612
	0755, 1507, 1545, 1568, 9526, 4819


1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-B*46 SSP subtypings.

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control band may sometimes be observed. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most tubes, or a band of 800 base pairs, for some tubes.

Tube number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-B*46 SSP subtyping.

In addition, tubes number 6 and 10 contain the primer pair giving rise to the shorter, 800 bp, internal positive control in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-B alleles one non-HLA-B*46 alleles will be amplified by primer mixes 1, 3, 5, 7 to 11, 13 and 14 In addition, the A*3304 allele will be amplified by primer mix 8.

4Specific PCR fragments shorter than 150 base pairs are less intense and not as sharp as longer specific bands.

5Primer mix 8 may give rise to a primer oligomer artifact.

[image: image1.wmf] Tube No. 

1

2

3

4

5

6

7

8

9

10

11

12

13

14

 HLA-B allele 

 *460101-460102 

1

5

9

 *4602 

1

2

5

9

 *4603 

1

3

5

 *4604 

4

5

9

 *4605 

1

9

10

 *4606 

1

5

6

 *4607N 

1

5

7

9

 *4608 

1

5

8

 *4609 

1

5

9

11

 *4610 

1

5

9

12

 *4611 

1

5

8

13

 *4612 

1

5

9

14

 *0709, 0711, 4038, 4052, 4059, 4209 

6

 *0755, 4819 

14

 *0815, 4073 

5



[image: image2.wmf]*1301, 1304, 1306, 1307N, 1310-1313, 1317, 1320, 1401-140202, 

1404-140602, 1502, 1503, 1505, 1506, 1509, 1510, 1513, 1516-1518, 

1520, 1521, 1523, 1525, 1529, 1531, 1536, 1537, 1539, 1540, 1544, 

1548, 1552, 1555, 1561, 1562, 1564, 1567, 1569, 1574, 1580, 1586, 

1588, 1589, 1591, 1593, 1595, 1598, 9503, 9506, 9508, 9512, 9514, 

9515, 9519, 9521, 9523, 9524, 9527, 9531-9534, 9536, 9539, 180101-

1812, 1814, 1815, 1817N, 1818, 1820, 1822-1827, 2701-270402, 

270502-2706, 2708-2710, 2712-2719, 2721, 2723, 2726-2732, 2735, 

2737, 2738, 350101, 350103-3507, 3510-3513, 3515-3517, 3519-

3530, 3532-3537, 3539-3542, 3546-3553N, 3555-3559, 3563, 3565Q, 

3568-3572, 3575-3578, 3582, 3583, 3702

8

380101-3817, 39010101-39010102L, 390103-3907, 3909, 3910, 3912-

3917, 3919, 3920, 3922-3932, 3934, 3935, 3937-3939, 3941-3943, 

4026, 4028, 4035, 4071, 4077, 440301-440303, 4407, 4410, 4413, 

4426, 4429, 4430, 4432, 4436-4440, 4443, 4445-4447, 4450, 4454, 

4457, 47010101-4705, 4802, 4812, 4904, 4905, 510101-5104, 5106, 

5107, 510901-510902, 5111N-5114, 5116-5119, 5121-5123, 5126-

5128, 5130-5135, 5137-5143, 5145, 5146, 5148-5151, 520101-5211, 

530101-5310, 5312, 5313, 5404, 5409, 5513, 5514, 5517, 5523, 

5527, 5528, 5602, 5604-5606, 5609-5612, 5615, 5618, 5621, 570101-

570103, 570301-570302, 5706-5708, 5710, 5711, 5714, 5801, 5802, 

5804-5815, 670101-6702, 7801-7805  

8

 *130201-1303, 1309, 1314-1316, 1319, 1542, 3560, 4901-4903, 

5001, 5002, 5004, 5115, 5401-5403, 5405N, 5407, 5408N, 5412, 

5413, 550101-550202, 5505, 5507, 5510-5512, 5515, 5516, 5518, 

5519, 5522, 5524-5526, 5601, 5607, 5608, 5614, 5616, 5617, 5901 

8

13

 *1308Q, 4415, 4418, 4501, 4503-4507, 5613, 5620 

13

 *1318, 5410, 5509 

3

13

 *1403, 1547, 1549, 1565, 1584, 1599, 1813, 350801-350803, 3545, 

3561, 3933, 4404, 4428, 4456N, 5105, 5129, 5311, 5702, 5704, 5705, 

5712 

3

 *1407N, 3940N,  

7

8

*15010101-150104, 150106, 1504, 1508, 151101-1512, 1514, 1515, 

1519, 1524, 1526N-1528, 1530, 1532, 1534, 1535, 1538, 1543, 1546, 

1550, 1553, 1554, 1556, 1558, 1560, 1563, 1566, 1570, 1571, 1573, 

1575-1577, 1579N, 1581, 1582, 1585, 1587, 1592, 1594N, 1596, 

1597, 9501, 9502, 9504, 9505, 9509-9511N, 9513, 9517, 9518, 9520, 

9522, 9528, 9529, 9535, 9537, 1819, 2725, 351401-351402, 3543, 

3544, 3567, 3936, 5406, 5603

9

 *1507, 1545, 1568, 9526 

9

14

 *1557 

w

w

9

 *9525 

9

10

 *9538, 4814 

6

8

 *3554 

8

10

 *3562 

9

11

 *3566 

3

6

 *3574 

5

8

 *3579 

w

 *3908, 3911 

w

8

 *3918 

8

9

 *5503 

5

8

13

 *5521 

9

13

 *5619N 

7

8

13

 A*3304 

8

Length of spec. PCR product

130

125

115

130

235

230

140

115

115

100

150

170

105

215

 Tube No. 

1

2

3

4

5

6

7

8

9

10

11

12

13

14

Specific Reaction

1st match: B*

2nd match: B*

Control Reaction

Failed Controls

False Positive

False Negative








Olerup SSP AB
In U.S. “For ReseSeptember Use Only.
www.olerup.com
Hasselstigen 1
Not For Use in Diagnostic Procedures.”
SE-133 33 Saltsjöbaden

Sweden

