# Gel Documentation Form and Worksheet

# HLA-B*53 (101.549-06/06u) 	Lot No: 8H0	Expiry Date: 2023-06-01

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

Interpretation:___________	                         Failed lanes: ___________                            Comments:_________________

Gel Picture

## PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 2, 7 and 22 may have tendencies of unspecific amplifications.

Primer mixes 11 and 14 may give rise to a lower yield of HLA-specific PCR product than the other HLA-B*53 primer mixes.

Primer mixes 3 and 19 have a tendency to giving rise to primer oligomer formation.

Primer mix 21 may give rise to a long fragment of approx. 600 bp in some HLA-B alleles. This band should not be considered in the interpretation of HLA-A*53 typings.

Primer mix 24 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-B*53 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.25.0, July 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3 The following HLA-B*53 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-B*53 alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- |
| 3 | 95bp 135 bp 290 bp 490 bp | *53:42 *53:03, 53:05:01-53:05:02, 53:16, 53:33 *53:15 *53:42 | *07:12, 07:18:01-07:18:02, 07:163, 07:199, 07:227, 07:248, 35:34, 35:39, 40:120, 44:107, 44:123, C*03:81, C*03:388, C*15:06:01-15:06:03, C*15:158 *18:54, 18:67, 18:108, 18:132, 27:52, 35:203, 35:355, 37:01:01:01-37:01:05, 37:01:07-37:05, 37:07-37:09, 37:12-37:13, 37:15-37:27, 37:30N-37:33N, 37:35-37:36, 37:38-37:60, 37:62-37:66, 37:68-37:76, 51:185, 52:09:01-52:09:02 *08:49, 08:60, 08:76, 08:129, 08:181, 41:48, 42:13 *07:65, 07:134, 35:34, 35:39, 35:199, 37:07, 37:09, 51:90, 51:135, 52:12 |
| 6 | 205 bp 240 bp | 53:02, 53:06, 53:28 *53:03, 53:05:01-53:05:02, 53:12, 53:16, 53:33 | *14:56, 15:38:01-15:38:02, 15:185, 15:335, 15:364, 15:368, 18:01:01:01-18:09, 18:12:01-18:15, 18:17N-18:20, 18:22-18:25, 18:27-18:34, 18:36-18:40, 18:42-18:55, 18:57:01-18:57:02, 18:59-18:60, 18:62-18:100, 18:102-18:109, 18:111-18:121, 18:123-18:147, 18:149-18:158, 35:21, 35:24:01-35:24:02, 35:188, 35:190, 35:287, 39:43, 51:37, 51:45, 51:63:01-51:63:02, 51:97, 51:216, 51:251, 52:39, 56:31, 57:14:01-57:14:02, 58:09, 58:76, C*02:142, C*06:207, C*07:516, C*07:521:01-07:521:02, C*12:87, C*14:76 *08:52, 08:117, 08:126, 13:25, 15:43, 15:87, 15:115, 27:106, 35:329, 37:55, 40:188, 40:394, 44:15:01:01-44:15:01:02, 44:109, 44:176, 44:257, 47:01:01:02-47:01:03, 47:05-47:10, 51:185, 52:09:01-52:09:02, 55:90, 56:46, 57:60 |
| 14 | 95 bp 130 bp 190 bp | *53:14 *53:10 | *07:09:01-07:09:02, 07:11, 07:17, 07:115, 07:237, 07:246, 07:287, 07:327, 08:28, 08:35, 08:37, 08:69, 08:200, 14:05, 15:07:01:01-15:07:03, 15:55, 15:68, 15:126, 15:197:02, 15:207, 15:331, 15:360, 15:383, 15:405, 15:421, 15:431, 15:450, 18:14, 35:05:01:01-35:05:03, 35:51, 35:58, 35:66, 35:72, 35:89, 35:97, 35:114, 35:232, 35:368, 35:389, 38:19, 39:03:01:01-39:03:01:02, 39:24:01-39:24:02, 39:37, 39:76, 39:120, 40:03:01:01-40:03:01:02, 40:105, 40:267, 40:284, 40:314, 41:18, 41:24, 41:51, 42:06, 42:09, 42:20, 46:12, 48:07, 48:14, 58:18, C*02:60, C*07:294, C*07:526:01, C*07:530, C*07:681, C*08:08:01-08:08:02 *35:349 *15:462, 57:01:01:01-57:01:22, 57:01:24-57:04:01, 57:04:03-57:09, 57:11-57:19, 57:21-57:35, 57:37-57:61, 57:63-57:88, 57:90-57:108, 57:110, 58:01:01:01-58:01:20, 58:01:21w, 58:01:22-58:02:02, 58:05-58:29, 58:31N-58:66, 58:68-58:73, 58:75-58:97 |
| 17 | 145 bp 275 bp | *53:30, 53:39 *53:20 | *40:13, 40:19, 40:109, 40:117, 40:292, 40:340, 40:394, 41:46, 44:18, 44:25, 44:50:01, 44:95, 49:01:01:01-49:01:13, 49:04:01-49:14, 49:16-49:25, 49:27-49:59, 51:112, 57:45, 57:51, 57:69 *51:187, A*25:37 |
| 18 | 110bp 140 bp | *53:48N *53:21 | *35:381N *07:134, 18:01:01:01-18:03:02, 18:05:01:01-18:15, 18:17N-18:24, 18:26-18:36, 18:38-18:47, 18:49-18:65, 18:67, 18:69-18:106, 18:108-18:124, 18:126-18:142, 18:144-18:158, 27:52, 35:50, 35:84, 35:162, 35:197, 35:231, 35:267, 37:01:01:01-37:05, 37:07-37:27, 37:29-37:30N, 37:32-37:41, 37:43-37:54, 37:56-37:60, 37:62-37:73, 37:75-37:76, 38:06-38:07, 39:19:01-39:19:02, 40:209, 40:350, 51:57, 51:103, 51:116, 51:218 |
| 22 | 115 bp 205 bp | *53:37 *53:19, 53:36 | *13:86, 35:02:01:01-35:02:07, 35:02:08w, 35:02:09-35:02:11, 35:81, 35:83, 35:95, 35:129N, 35:146, 35:154, 35:162, 35:172, 35:182-35:184, 35:211, 35:220, 35:258, 35:270, 35:273, 35:285, 35:311, 35:323, 35:335, 35:339, 35:357, 35:361, 35:366-35:367, 35:372, 35:374, 35:377-35:379, 35:387, 35:391, 35:394, 35:396, 50:36, 58:65 |

‘w’, may be weakly amplified.

Changes in revision R01 compared to R00:

1. The internal positive control band length in primer mix 22 was corrected from 800 bp to 1070 bp.

Changes in revision R02 compared to R01:

1. The expiration date has been altered due to shelf-life extension.