Gel Documentation Form and Worksheet

HLA-B*57:01 (101.572-12/12u) 	       Lot No: 8L3	Expiry Date: 2025-01-01

Name:___________________                                Test Date: ____________ 		Review Date:_______________

Sample ID:_______________                                 Tested By: _____________		Reviewed By:______________

DNA Conc.(ng/ul):_________

Interpretation:___________		      Failed lanes: ___________	   Comments:_________________

Gel Picture

PHOTO DOCUMENT

Abbreviations

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 2 may give rise to a lower yield of HLA-specific PCR product than the other B*57:01 primer mixes.

Primer mix 18 has tendency of unspecific amplifications.

Primer mix 24 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR products generated by the HGH positive control primer pair are 200 base pairs.

1HLA-B alleles listed on the IMGT/HLA web page 2020-10-15, release 3.42.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3 The following HLA-B*57 primer mixes have two or more product sizes:

| Primer Mix |  | Size of spec. PCR product | Amplified HLA-B*57:01 alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- | --- |
| 3 | 95 bp 170 bp 215 bp |  |  | *57:04:01-57:04:03, 57:32, 44:153 *57:06, 57:18 *57:27 |
| 4 | 90 bp 105 bp 170 bp |  |  | *57:20, 57:62 *57:15, 57:29, 14:61, 15:214, 18:81, 35:250, 37:93, 39:145, 39:158, 40:427, 44:311, 51:165, 55:108 *57:91, 53:59 |
| 5 | 90 bp 165 bp 200 bp 245 bp |  |  | *57:16, 57:34, 58:54 *57:23 *57:46, 57:64 *57:07, 57:26, 57:93, 57:119, 08:181, 44:153, 55:14, C*06:72 |
| 6 | 90 bp 205 bp |  |  | *57:02:01:01-57:03:05, 57:07, 57:09, 57:12, 57:17, 57:39, 57:42, 57:46, 57:57, 57:63, 57:66, 57:70, 57:80, 57:84, 57:94, 57:96, 57:101, 57:124, 57:137, 08:181, 40:30, 40:34 *57:08, 57:65, 15:403 |
| 7 | 100 bp 140 bp 185 bp 215 bp 250 bp |  |  | *57:53 *57:13, 57:31:01-57:31:02, 07:227, 08:181, 40:30, 40:34, C*06:72 *57:50 *57:09, 57:24, 57:119, 08:181 *57:13, 57:14:01-57:14:02, 57:25, 07:227, 40:30, 40:34 |
| 8 | 160 bp 210 bp |  |  | *57:87 *57:10, 57:39, 57:51, 57:73, 07:219, 15:446, 58:41 |
| 9 | 110 bp 150 bp |  |  | *57:33, 44:450 *57:21, 57:40, 57:74, 14:20, 15:461, 27:188, 35:127, 44:317, 50:54, 51:186, 53:41, 55:100 |
| 10 | 205 bp 240 bp |  |  | *57:13, 57:22, 57:57, 57:63, 07:227, 55:14, C*06:72w *57:43 |
| 12 | 75 bp 140 bp |  |  | *57:35 *57:08, 57:60, 15:87, 15:403 |
| 13 | 105 bp 135 bp 200 bp |  |  | *57:45, 57:51, 57:69, 57:71, 57:88, 58:111, A*02:285, A*02:799 *57:58, 57:80, 57:90, 07:219, 58:12, A*02:42:01-02:42:02, A*02:310, A*02:597 *57:49 |
| 14 | 115 bp 135 bp 205 bp |  |  | *57:54, 57:96, 15:33, 15:248, 49:22, 51:126, C*03:87:01-03:87:02, C*03:414, C*05:27, C*05:39, C*05:184, C*08:115, C*08:182 *57:47 *57:77, 44:36, C*03:232 |
| 15 | 145 bp 185 bp 295 bp |  |  | *57:56, 57:78, 15:340, 35:300, 40:218, 44:307, 45:21, 46:28, 51:58, 52:82, 58:83 *57:48, A*23:72, A*33:12 *57:55, 57:94 |
| 17 | 90 bp 180 bp |  |  | *57:04:01-57:04:03, 57:41, 44:153, C*06:72 *57:68 |
| 18 | 170 bp 275 bp 300 bp |  |  | *57:37, 57:79N *57:81, 44:13, 44:67, 51:170 *57:86, 53:25 |
| 19 | 100 bp 165 bp 215 bp |  |  | *57:76 *57:92 *57:72 |
| 20 | 240 bp 340 bp |  |  | *57:44, 18:167, 44:249 *57:75, 35:374, 40:327?, 52:81 |
| 21 | 100 bp 135 bp 485 bp |  |  | *57:36 *57:38 *57:05 |

Abbreviations

‘w’, might be weakly amplified.

‘?’, nucleotide sequence of the primer matching region not known.