Gel Documentation Form and Worksheet

HLA-B*67 (101.550-06/06u)	                                      Lot: 7E9                                        Expiry Date:   2019-05-01

Name:____________________

Sample ID:________________

DNA Conc.(ng/ul):___________

Test  Date: _______________

Tested By: _______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________	Failed lanes: ___________ 	 Comments:_________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

Notes: Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 1 and 5 may give rise to a lower yield of HLA-specific PCR product than the other B*67 primer mixes.

Primer mix 8 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-B*67 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.25.0, July 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3Primer mix 7: Specific PCR fragment of 85 bp in the B*67:04 and the B*07:94, 08:106 and 40:115 and in the C*03:124, C*07:262, 07:417, 07:475 and C*12:95 alleles. Specific PCR fragment of 165 bp in the B*67:05 and the B*15:20, 15:228, 15:305, 35:01:01:01-35:01:27, 35:01:29-35:30, 35:32:01-35:34, 35:36-35:42:02, 35:44-35:45, 35:47-35:56, 35:58-35:59:02, 35:61:01-35:72, 35:74-35:78, 35:80-35:86, 35:88-35:184, 35:186-35:278, 35:280-35:322, 48:02:01-48:02:03, 51:01:01:01-51:24:05, 51:26-51:41N, 51:43-51:46, 51:48-51:156, 51:158:01-51:163, 51:165-51:208, 52:01:01:01-52:06:03, 52:08-52:61, 53:01:01-53:43, 56:05:01-56:05:02, 56:21, 56:36, 58:01:01:01-58:02:02, 58:04-58:10N, 58:12-58:19, 58:21-58:29, 58:31N-58:81, 78:01:01-78:09, 81:03 and 83:01?

4The following HLA-B*67 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| B*67:04, 67:05 | 7 |

The HLA-B*67 subtyping kit cannot distinguish the silent mutations in the B*67:01:01-67:01:03 alleles.

‘w’, may be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.