HLA-B*67 (101.550-06/06u)	                                      Lot: 6S0                                        Expiry Date:   2027-10-01

Name:____________________

Sample ID:________________

DNA Conc.(ng/ul):___________

Test Date: _______________

Tested By: _______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________	Failed lanes: ___________ 	 Comments:_________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

Notes: Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 1 may give rise to a lower yield of HLA-specific PCR product than the other B*67 primer mixes.

Primer mix 3 may have a tendency to giving rise to primer oligomer formation.

Primer mix 7 may have tendency of unspecific amplification.

Primer mix 8 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1HLA-B*67 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.35.0, January 2019.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3 The following HLA-B*67 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-B*67 alleles | Other amplified HLA-B alleles |
| --- | --- | --- | --- |
| 7 | 85 bp 165 bp | *67:04 *67:05 | *07:94, 08:106, 40:115, 55:99, C*01:166, C*03:124, C*07:262, C*07:417, C*07:475, C*08:153, C*12:95 *15:20, 15:228, 15:305, 15:435, 15:599, 15:658, 27:197, 35:01:01:01-35:01:27, 35:01:29-35:01:46, 35:01:48-35:30:01:02, 35:32:01-35:32:02, 35:33-35:34, 35:36-35:42:03, 35:45, 35:47-35:56, 35:58-35:59:02, 35:61:01-35:72, 35:74-35:78, 35:80-35:86, 35:88-35:134N, 35:136-35:184, 35:186-35:208:01, 35:209-35:278, 35:280-35:358, 35:360-35:466, 35:468-35:470, 35:472-35:474, 35:477-35:482, 35:484:01-35:533, 35:535-35:544, 35:546-35:548, 35:550-35:586, 37:80, 48:02:01-48:02:03, 49:66, 51:01:01:01-51:01:55, 51:01:57-51:24:05, 51:26-51:41N, 51:43-51:46, 51:48-51:156, 51:158:01-51:163, 51:165-51:219, 51:221-51:236, 51:238-51:294, 51:296-51:299, 51:301-51:306N, 51:308-51:358, 51:360-51:386, 52:01:01:01-52:01:50, 52:01:52-52:06:03, 52:08-52:107, 52:109-52:115N, 53:01:01:01-53:01:27, 53:01:29-53:58, 53:60-53:73, 56:05:01, 56:21, 56:36, 56:94, 58:01:01:01-58:01:21, 58:01:23-58:01:44, 58:02:01:01-58:02:02, 58:04-58:10N, 58:12-58:19, 58:21-58:29, 58:31N-58:104, 58:107-58:143, 78:01:01:02-78:10, 81:03 |

4The following HLA-B*67 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix |
| --- | --- |
| B*67:04, 67:05 | 7 |

‘w’, may be weakly amplified.