Gel Documentation Form and Worksheet

HLA-B*81 (101.553-06/06u)	                                       Lot: 1H5                                     Expiry Date: 2021-05-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):________

Test Date: ____________

Tested By: ____________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                             Failed lanes: ____________                            Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon size

Notes: Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 3 may have tendency of unspecific amplification.

Primer mix 3 may have tendency of giving rise to primer oligomer formation.

Primer mix 10 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-B*81 alleles in bold lettering are listed as confirmed alleles on the on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.33.0, July 2018.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-B*81 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-B*81 alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- |
| 5 | 105 bp 165 bp | *81:05 *81:03 | *07:83 *15:20, 15:228, 15:305, 15:435, 35:01:01:01-35:01:27, 35:01:29-35:01:46, 35:01:48-35:30, 35:32:01-35:34, 35:36-35:42:02, 35:45, 35:47-35:56, 35:58-35:59:02, 35:61:01-35:72, 35:74-35:78, 35:80-35:86, 35:88-35:134N, 35:136-35:184, 35:186-35:278, 35:280-35:358, 35:360-35:372, 48:02:01-48:02:03, 51:01:01:01-51:01:55, 51:01:57-51:24:05, 51:26-51:41N, 51:43-51:46, 51:48-51:156, 51:158:01-51:163, 51:165-51:219, 51:221-51:236, 52:01:01:01-52:06:03, 52:08-52:76, 53:01:01-53:50, 56:05:01-56:05:02, 56:21, 56:36, 58:01:01:01-58:01:21, 58:01:23-58:02:02, 58:04-58:10N, 58:12-58:19, 58:21-58:29, 58:31N-58:93N, 67:05, 78:01:01:01, 78:01:02-78:09 |
| 6 | 80 bp 105 bp | *81:04N *81:05 | *07:83 |

4The B*81:02 and the B*07:202 and B*07:279 alleles will give rise to identical amplification patterns. These alleles can e.g. be distinguished by the HLA-B low resolution kit and/or HLA-B*07 high resolution kit.

‘w’, may be weakly amplified.

‘?’, the nucleotide sequence of the primer matching region is not known.