# Gel Documentation Form and Worksheet

# HLA-C high resolution for frequent alleles

# (101.707-12/12u) 	Lot No: 1F9	Expiry Date: 2021-06-01

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________	_

Interpretation:_______________	                    Failed lanes: ____________                           Comments:_____________

Gel Picture

## PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 3, 5, 18, 21, 40, 44, 45, 57, 61, 62 and 86 may have tendencies of unspecific amplifications.

Primer mixes 6, 20, 36, 78 and 80 have a tendency to giving rise to primer oligomer formation.

Primer mixes 10, 12, 17, 27, 53, 64, 72, 77, 82, 86 and 87 may give rise to a lower yield of HLA-specific PCR product than the other HLA-C high resolution primer mixes.

Primer mixes 10, 44 and 92 may give rise to a long unspecific amplification product. This band should be disregarded when interpreting the typing results.

Primer mix 64 may give rise to unspecific amplification in C*12 alleles.

Primer mix 19 may give rise to a lower yield of HLA-specific PCR product than the other HLA-C high resolution primer mixes for the C*03 alleles.

In primer mix 8 the positive control band may be weaker than for other HLA-C high resolution primer mixes.

Primer mix 96 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

Change in revision R01 compared to R00:

- In primer mix 8 the positive control band may be weaker than for the other HLA-C*high resolution primer mixes. A footnote has been added in the Specificity Table.

Change in revision R02 compared to R01:

1.	The expiration date has been altered due to extension of shelf-life.

1HLA-C alleles listed on the IMGT/HLA web page 2016-October-14, release 3.26.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-C*high primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-C alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- |
| 2 | 100 bp 215 bp | *04:05, 07:54 *01:05, 01:77, 04:240, 05:111, 07:37, 07:216, 07:307, 14:08, 16:06 | B*07:77, B*07:193 |
| 3 | 65 bp 100 bp 200 bp 225 bp 290 bp 365 bp | *08:99 *01:27 *01:06 *07:06, 07:18-07:19, 07:330:02, 07:517 *01:11 *15:13:01:01, 15:103 |  |
| 4 | 100 bp 200 bp | *02:34, 16:09 *01:07:01-01:07:02, 06:10, 06:43:01, 14:24:02 |  |
| 5 | 130 bp 155 bp 215 bp 275 bp 300 bp | *02:02:28, 03:03:29, 03:04:19, 03:113:01, 04:01:43, 05:01:17, 05:09:02, 05:17, 06:02:23, 06:27, 07:03, 07:17:01-07:17:03, 07:140-07:141:02, 07:304, 07:307, 07:512, 08:15:02, 08:72:02, 12:03:10, 16:02:11, 16:47 *07:23, 12:13 *01:08, 12:19 *06:27, 16:47 *06:02:23, 07:31:01-07:31:02, 07:177, 07:514, 12:03:10, 16:02:11 | A*26:55, A*29:27, B*07:02:43, B*07:33:03, B*08:01:25, B*15:17:04, B*18:90, B*35:01:38, B*35:03:15, B*35:08:04, B*35:15:02, B*37:04:02, B*39:09:03, B*57:01:18, B*58:01:17 |
| 6 | 105 bp 150 bp 190 bp 210 bp 265 bp | *03:24, 15:90 *01:20 *01:60, 07:05, 07:177, 07:364, 07:538 *01:04, 01:09, 06:23, 06:179, 12:178 *01:15, 02:114, 03:33 | B*27:36, B*48:21, B*48:26 B*15:129, B*15:342, B*27:25, B*27:115 |
| 8 | 100 bp 135 bp 595 bp | *02:03, 02:16:02, 02:18, 03:03:22, 03:04:34, 06:07, 07:02:75 *02:20 *07:179 | B*27:34, B*35:01:30w, B*40:02:07, B*40:06:02, B*51:01:52w, B*57:01:16w, B*57:03:02w |
| 10 | 105 bp 130 bp | *03:342, 07:46, 07:356 *01:115, 02:07, 02:97, 14:79, 16:10, 16:67 | B*56:08, 56:14 |
| 11 | 75 bp 160 bp 505 bp | *02:08, 02:42, 03:11:01-03:11:02, 03:18:02, 03:37:01-03:37:02, 03:64:01, 03:301, 04:256, 06:03:01, 07:314:02, 12:03:23, 15:10:02-15:10:03, 15:113 *02:104, 03:46, 03:248, 04:16, 15:100 *07:06, 07:330:02, 07:343 | B*58:74 |
| 12 | 70 bp 280 bp | *07:53, 07:216 *02:02:01-02:02:03, 02:02:05-02:02:29, 02:02:31-02:13, 02:15-02:25Q, 02:26:02-02:32, 02:34-02:40:02, 02:42-02:86, 02:88-02:106, 02:108-02:126, 03:02:01-03:02:09, 03:02:11-03:02:15, 03:04:01:01-03:04:25, 03:04:27-03:10, 03:14-03:17:01, 03:19, 03:23-03:29, 03:32-03:38:02, 03:40:01-03:42, 03:44-03:48, 03:51, 03:54, 03:57:01-03:57:02, 03:60, 03:63-03:64:01, 03:65, 03:70-03:74, 03:77-03:78, 03:80:01-03:80:02, 03:82, 03:84, 03:87:01-03:87:02, 03:89-03:95, 03:98, 03:100-03:101, 03:104-03:111, 03:113:01-03:115, 03:117-03:118, 03:121N, 03:123, 03:125, 03:128-03:131, 03:134-03:136, 03:138-03:140, 03:143, 03:145-03:149, 03:153-03:155, 03:157, 03:159, 03:162-03:164, 03:169Q, 03:172-03:174, 03:178-03:181, 03:183-03:184:02, 03:186:01-03:186:02, 03:190-03:191, 03:193-03:194, 03:197-03:201N, 03:208N-03:213, 03:215-03:216, 03:218-03:219, 03:221-03:222, 03:224N-03:226, 03:232-03:236, 03:238-03:240, 03:244Q-03:250, 03:252, 03:255-03:261, 03:263:01-03:263:02, 03:265N-03:266, 03:269-03:270, 03:277N-03:283, 03:286-03:287, 03:292-03:294, 03:296-03:303, 03:305-03:306, 03:309-03:311, 03:313-03:315, 03:317-03:318N, 03:322-03:323N, 03:326, 03:328-03:334, 03:337-03:338, 03:340, 03:342-03:343, 04:03:01-04:03:03, 04:06, 04:16, 04:80, 04:107, 04:147, 04:160, 04:171, 04:256, 05:58:01, 05:58:03, 06:03:01, 06:132:01-06:132:02, 07:96:01-07:96:02, 07:314:02, 12:03:23, 15:02:01:01-15:02:20, 15:02:22-15:02:24, 15:02:26-15:05:08, 15:05:10-15:09, 15:10:02-15:11, 15:13:01:01, 15:15-15:19, 15:21-15:22, 15:24-15:35, 15:37-15:60, 15:62, 15:64-15:106, 15:108-15:133, 16:34, 16:70 | B*40:164, B*51:01:28, B*58:74 |
| 16 | 110 bp 240 bp 275 bp 370 bp | *07:104N, 12:46N *03:53, 12:07, 12:17 *01:65, 03:05, 03:13:01-03:13:02, 03:17:01, 03:25, 03:27, 03:35:01-03:35:02, 03:71, 03:167, 03:178, 03:198, 03:249, 03:267, 03:271, 03:292, 03:296, 03:302, 04:212, 05:86, 07:133, 07:242, 07:330:01 *15:22, 15:37, 15:55, 15:58, 15:65, 15:72, 15:102, 15:125 | B*15:272N, B*40:286N B*35:310 B*15:96 |
| 17 | 110 bp 180 bp 565 bp | *01:118, 01:128, 03:08, 03:29, 03:31, 03:246, 06:44, 07:01:01:01-07:01:24, 07:01:26-07:01:58, 07:06-07:07, 07:09, 07:16, 07:18-07:20, 07:22, 07:24, 07:26:01-07:26:02, 07:28, 07:30, 07:35-07:36, 07:40, 07:44, 07:52-07:53, 07:55N, 07:57-07:60, 07:65, 07:70-07:71, 07:73, 07:77-07:78:02, 07:81-07:83, 07:86, 07:89, 07:91-07:96:02, 07:98N, 07:103-07:104N, 07:106, 07:108:01-07:113, 07:115-07:116, 07:118-07:120, 07:122, 07:124, 07:128-07:129, 07:131:01-07:132, 07:134, 07:140-07:141:02, 07:148, 07:150Q-07:151, 07:153, 07:162, 07:164N-07:166, 07:170, 07:173, 07:176, 07:179-07:180, 07:182, 07:188-07:191N, 07:196-07:197, 07:200-07:201, 07:203-07:207, 07:210, 07:212, 07:214-07:215, 07:219, 07:222-07:224, 07:227N-07:228, 07:230-07:231, 07:235Q-07:237, 07:246:01-07:250, 07:253-07:257:04, 07:263, 07:266-07:269, 07:276-07:282, 07:292-07:293, 07:296-07:301, 07:303-07:304, 07:310-07:311, 07:313, 07:317, 07:330:01-07:332, 07:337, 07:343, 07:353, 07:362, 07:366, 07:369-07:371, 07:377, 07:385-07:387, 07:404, 07:407-07:408, 07:410, 07:412, 07:417, 07:419, 07:421-07:422, 07:424, 07:427, 07:430-07:431, 07:435, 07:437N, 07:440-07:444, 07:448, 07:450, 07:453-07:455, 07:458, 07:460-07:463, 07:468-07:471, 07:475-07:476N, 07:479, 07:481, 07:488, 07:490-07:492, 07:498-07:499, 07:502, 07:506, 07:508, 07:512, 07:514-07:515, 07:517, 07:521-07:522, 07:526-07:527, 07:531, 07:536-07:537, 07:540, 07:545-07:547, 07:549-07:550, 07:554, 07:555w, 07:557, 07:559, 07:561, 14:73, 15:02:10, 15:02:17, 15:99, 18:05 *01:51, 03:03:01:01-03:03:25, 03:03:27-03:03:34, 03:11:01-03:11:02, 03:13:01-03:13:02, 03:18:01-03:18:02, 03:20N-03:22Q, 03:30-03:31, 03:39, 03:43:01-03:43:02, 03:49-03:50, 03:52-03:53, 03:55-03:56, 03:58-03:59, 03:61-03:62, 03:66-03:69, 03:75-03:76, 03:79, 03:81, 03:83, 03:85-03:86, 03:88, 03:96-03:97, 03:102-03:103, 03:112, 03:116:01-03:116:02, 03:119:01-03:120, 03:122, 03:124, 03:126-03:127, 03:132-03:133, 03:141-03:142, 03:144, 03:150-03:152, 03:158, 03:160-03:161, 03:165, 03:167-03:168, 03:171, 03:175-03:177, 03:182, 03:185, 03:187-03:189N, 03:192, 03:195-03:196, 03:202-03:207, 03:214, 03:217, 03:220, 03:223, 03:227-03:231, 03:237, 03:241-03:243, 03:251, 03:253-03:254, 03:262, 03:267-03:268, 03:271-03:276, 03:284-03:285, 03:288-03:291, 03:295, 03:304, 03:307-03:308, 03:312, 03:316N, 03:319-03:321, 03:324-03:325, 03:327, 03:336, 03:339, 03:341, 06:53:01-06:53:02, 07:56:01-07:56:02, 07:60, 07:79, 07:109, 14:46:02, 17:26 *01:02:18, 06:02:38, 07:447, 12:03:04, 12:03:09, 12:195:01, 14:02:08 | B*07:259 B*07:93, B*08:39, B*27:155, B*35:101:02, B*35:195, B*35:211, B*38:27, B*51:76, B*54:34 B*27:05:27 |
| 18 | 90 bp 120 bp 185 bp 225 bp 255 bp | *01:65, 03:21, 03:28, 03:49, 03:80:02, 03:117, 03:142, 03:287, 05:86 *03:14, 03:162, 03:220, 03:261, 07:133, 07:330:01 *07:21 *03:23 *03:09, 03:22Q, 03:42, 03:71, 03:89, 03:116:01, 03:119:01-03:119:02, 03:139, 03:169Q, 03:199, 03:240, 03:249, 03:267, 03:292, 03:300, 04:212, 05:86, 07:36, 07:133, 07:242, 07:330:01, 07:355 | B*15:96 B*56:37 |
| 19 | 110 bp 265 bp 305 bp 540 bp | *03:14, 03:36, 03:77, 03:162, 03:187, 03:220, 03:261, 03:296, 04:212, 07:133, 07:330:01 *03:13:01-03:13:02, 03:35:01-03:35:02, 03:178, 03:198, 03:267, 03:292, 04:212, 05:86, 07:133, 07:242, 07:330:01 *03:32, 03:45, 03:136, 03:297, 04:80, 04:100, 04:178, 06:14, 06:143, 07:10, 07:43:01-07:43:02, 07:184, 07:196, 15:25, 15:62 *03:15, 03:32, 03:45, 03:60, 03:136, 03:297, 04:212, 07:133, 07:242, 07:330:01 | B*15:96 B*15:357, B*35:178, B*35:282, B*35:316 |
| 20 | 240 bp 265 bp | *02:02:01-02:02:03, 02:02:05-02:02:29, 02:02:31-02:13, 02:15-02:25Q, 02:26:02-02:26:03, 02:28-02:40:02, 02:42-02:64, 02:66-02:86, 02:88-02:106, 02:108-02:114, 02:116-02:126, 03:07:01, 03:10, 03:15, 03:29, 03:45, 03:163, 03:268, 03:297, 04:03:01-04:03:03, 04:06, 04:16, 04:80, 04:107, 04:147, 04:160, 04:171, 04:256, 05:13, 05:58:01, 05:58:03, 06:03:01, 06:132:01-06:132:02, 15:02:01:01-15:02:20, 15:02:22-15:02:24, 15:02:26-15:05:08, 15:05:10-15:06:03, 15:08-15:09, 15:10:02-15:13:01:01, 15:15-15:19, 15:21-15:22, 15:24, 15:26-15:35, 15:37-15:42, 15:44:01-15:60, 15:62, 15:64-15:84Q, 15:86-15:106, 15:108-15:133, 16:70 *01:14, 15:104 |  |
| 21 | 195 bp 225 bp 250 bp 280 bp | *03:30, 03:65, 04:18, 04:52, 04:52, 04:55, 05:55, 12:58, 14:10, 15:12 *04:19 *03:50 *03:251, 03:314, 04:120, 05:10, 05:64:01-05:64:02, 08:19, 08:44, 08:61, 08:82, 08:101, 08:126, 08:143, 15:130 | B*44:148, B*58:76 |
| 22 | 80 bp 140 bp | *01:97, 03:02:01-03:02:15, 03:15-03:16, 03:33, 03:36, 03:40:01, 03:40:03-03:40:04, 03:42-03:43:01, 03:60, 03:71, 03:84, 03:89, 03:95, 03:108, 03:110, 03:119:02, 03:121N, 03:132, 03:139, 03:146, 03:169Q, 03:175, 03:194, 03:197-03:201N, 03:216, 03:221-03:222, 03:224N-03:226, 03:238, 03:240, 03:245, 03:248, 03:258, 03:264, 03:271, 03:279, 03:298-03:301, 03:315, 03:329-03:330, 03:332, 03:338, 04:54, 06:02:08, 07:02:32, 07:17:03, 07:242, 12:03:20, 14:02:01:01-14:02:01:02, 14:02:03-14:16, 14:18-14:24:01, 14:25, 14:27-14:53, 14:56-14:78, 14:80-14:81, 15:77, 16:01:06 *01:97, 03:17:01, 03:71, 03:249, 03:271, 03:302, 04:37, 04:230, 06:02:08, 06:34:02, 12:03:20, 14:02:01:01-14:02:01:02, 14:02:03-14:07N, 14:10-14:16, 14:18-14:27, 14:29-14:53, 14:55-14:81, 16:01:06 | B*15:78:03 |
| 23 | 150 bp 240 bp | *01:14, 02:02:13, 02:02:29, 03:07:01, 03:10, 03:15, 03:29, 03:45, 03:163, 03:268, 03:297, 04:01:23, 05:01:12, 05:29:02, 06:02:01:01-06:02:01:03, 06:02:03-06:04:02, 06:06-06:10, 06:12-06:76:01, 06:77-06:81, 06:83-06:146, 06:148-06:196, 07:07, 07:09, 07:76:01-07:76:02, 07:315, 07:328, 07:406, 07:559, 12:04:01, 14:12, 14:49, 15:02:10, 15:02:17, 15:99, 17:01:01:02-17:01:03, 17:01:05-17:21, 17:23-17:34, 18:01-18:10 *05:13 | B*15:277, B*35:222, B*42:22 |
| 26 | 145 bp 170 bp | *04:07, 04:27, 04:89, 04:153, 05:78:01-05:78:02, 07:64, 07:402, 15:36 *02:09 |  |
| 27 | 135 bp 175 bp 220 bp | *01:35, 01:107, 04:08, 04:34, 04:147, 04:212, 14:20, 18:08 *04:30, 15:15, 15:77 *04:09N |  |
| 29 | 130 bp 190 bp 290 bp | *03:05, 03:13:02, 03:25, 03:27, 03:35:02, 03:135, 03:167, 03:296, 04:15:01, 04:15:03, 14:09 *02:15 *07:14 | B*15:78:03 B*07:221 |
| 31 | 150 bp 285 bp | *01:13, 02:51, 03:87:01-03:87:02, 04:223:01-04:223:02, 05:09:01-05:09:03, 05:17, 05:52, 07:130, 08:15:01-08:15:02, 08:51, 08:73, 12:12, 12:144, 12:185, 16:27 *02:57, 04:40, 04:242, 06:123, 07:45 | B*15:33, B*15:248, B*35:08:02, B*35:08:05 B*38:23, B*39:73 |
| 32 | 120 bp 180 bp 210 bp | *05:03, 07:52 *03:48 *16:07:01 |  |
| 33 | 190 bp 225 bp | *04:28 *05:04:01-05:04:02, 05:103:01-05:103:02, 05:135, 06:129, 07:68, 07:260:01-07:260:02, 07:302, 08:09, 08:11, 08:59, 08:113 | B*15:337, B*18:83, B*58:76 |
| 34 | 220 bp 275 bp | *03:231w, 04:10, 04:11w, 04:36w, 04:55w, 04:153, 04:169, 04:210, 04:214w, 04:215N, 05:78:01-05:78:02, 15:36 *03:251, 03:314, 04:27, 05:05:01-05:05:02, 05:07N, 05:135, 08:08:01-08:08:02, 08:62, 08:82 | A*02:425, A*02:519, A*29:10:01-29:10:02, A*68:69, B*14:32, B*15:337, B*18:83, B*44:148 |
| 36 | 85 bp 120 bp 160 bp 210 bp 235 bp | *01:65, 03:28, 03:49, 03:117, 05:06, 05:86 *05:15 *03:25, 04:129, 05:42, 05:46, 06:67, 07:101, 07:148, 07:161, 08:05, 08:21, 08:137, 12:127, 17:05 *03:06:01-03:06:02, 03:26 *03:88 |  |
| 37 | 80 bp 130 bp 195 bp 235 bp 280 bp | *07:50 *05:23, 05:62, 05:134, 07:04:01:01-07:04:11, 07:11, 07:63, 07:68, 07:101, 07:139, 07:142, 07:181, 07:199:01-07:199:02, 07:260:01-07:260:02, 07:272, 07:302, 07:323-07:324, 07:328-07:329N, 07:336, 07:338, 07:354-07:355, 07:357-07:358, 07:361, 07:364-07:365, 07:378, 07:394-07:395, 07:403, 07:406, 07:420, 07:426, 07:428, 07:459, 07:466-07:467, 07:480, 07:487, 07:501, 07:523, 07:534-07:535, 07:552, 07:562-07:563, 08:07, 08:47, 08:104 *05:14 *05:19 *05:16, 05:85, 05:107, 06:129, 07:364, 08:12 | B*14:32 |
| 43 | 100 bp 155 bp 240 bp 275 bp | *01:02:34, 01:21, 02:42, 02:107, 04:140, 04:166, 04:220, 05:98, 06:05, 07:02:09, 08:14, 08:80, 08:103, 12:16, 12:147, 15:63, 15:113, 16:80 *06:13 *01:10, 01:51, 02:05:01-02:05:03, 02:17, 03:11:01-03:11:02, 03:39, 06:08, 06:53:01-06:53:02, 07:56:01-07:56:02, 07:60, 07:79, 07:109, 12:119, 14:25, 16:29, 16:50, 17:21 *06:27, 16:47 | B*67:02 B*07:93, B*07:239, B*08:39, B*14:46, B*38:27, B*40:243 |
| 44 | 90 bp 135 bp 165 bp | *03:20N *16:08 *02:22, 02:118, 04:14, 04:68, 04:94:01-04:94:02, 05:08, 05:52, 05:89, 05:106:01-05:106:02, 05:112, 06:09, 06:144, 08:27, 08:29, 08:31, 12:31, 12:144, 16:45, 17:04, 18:03 | A*24:96, 24:146, B*15:137, B*50:47 |
| 45 | 140 bp 255 bp 550 bp | *06:11, 06:122w, 06:124w, 06:147, 07:01:13, 07:04:01:01-07:04:11, 07:11-07:12, 07:45, 07:63, 07:68, 07:101, 07:139, 07:142, 07:181, 07:199:01-07:199:02, 07:272, 07:302, 07:323-07:324, 07:329N, 07:338, 07:354-07:355, 07:358, 07:364-07:365, 07:378, 07:394-07:395, 07:420, 07:426, 07:428, 07:447, 07:459, 07:466-07:467, 07:480, 07:487, 07:501, 07:523, 07:534-07:535, 07:552, 07:562-07:563, 12:02:11, 12:03:09, 12:11, 12:118, 16:01:16 *03:47, 12:126 *06:02:38, 07:447, 12:03:09 |  |
| 50 | 240 bp 305 bp 415 bp | *03:03:29, 03:04:19, 03:113:01, 07:29:01-07:29:02, 07:177, 07:184, 07:248 *03:03:29, 03:04:19, 03:113:01, 07:10, 07:43:01-07:43:02, 07:184, 07:196, 07:367 *06:17, 07:07, 07:09, 07:49, 07:61N, 07:76:01-07:76:02, 07:210, 07:238, 07:247, 07:315, 07:328, 07:403, 07:406, 18:01-18:10 | B*35:01:38, B*35:03:15, B*35:08:04, B*35:15:02, B*37:04:02, B*57:01:18, B*58:01:17 |
| 51 | 100 bp 235 bp 265 bp | *07:32N, 07:67 * 01:16- 01:18, 01:21, 01:23, 01:128, 07:316, 07:338 * 01:19, 07:24, 07:218 |  |
| 53 | 80 bp 115 bp 265 bp | *07:11 *05:12 *05:29:01-05:29:02, 08:13, 08:16:01, 08:25, 08:94 |  |
| 55 | 165 bp 215 bp 245 bp 320 bp 345 bp 405 bp 465 bp | *02:104, 03:248, 04:11, 04:33, 04:169, 04:172, 05:141, 12:194, 15:100, 16:62 *07:26:01-07:26:02, 07:33N, 07:92, 07:96:01-07:96:02, 07:314:01-07:314:03, 07:317, 07:351, 18:09 *07:03 *07:72 *01:12:01-01:12:02, 01:21, 01:34, 01:49:01-01:50, 01:79:01-01:79:02, 01:101-01:102 *07:61N *07:07, 07:16, 07:51, 07:177, 07:181, 07:328, 07:367 | B*07:267 |
| 56 | 100 bp 130 bp 250 bp 305 bp | *07:32N *02:02:28, 06:02:23, 07:22, 07:25, 07:137:01, 07:138, 07:404, 12:03:10 *07:35, 07:352 *03:03:29, 03:04:19, 03:113:01, 07:10, 07:43:01-07:43:02, 07:184, 07:196, 07:367 | A*02:01:126, A*03:01:46, A*03:154:02, A*11:32:02, A*24:02:62, A*24:347:01, A*25:27:02, A*33:03:09, A*68:01:12 B*35:01:38, B*35:03:15, B*35:08:04, B*35:15:02, B*37:04:02, B*57:01:18, B*58:01:17 |
| 65 | 150 bp 415 bp | *12:06 *01:118, 01:128, 03:08, 03:29, 03:31, 03:246, 04:112, 04:169, 05:36, 06:44, 12:08, 12:81, 12:188, 14:73, 16:64, 16:70, 16:87 |  |
| 71 | 130 bp 200 bp | *03:02:01-03:02:07, 03:02:09-03:03:31, 03:03:33-03:04:26, 03:04:28-03:04:29, 03:04:31-03:04:54, 03:06:01-03:11:02, 03:13:01-03:13:02, 03:15-03:24, 03:26, 03:28-03:40:01, 03:40:03-03:43:01, 03:44-03:57:02, 03:59-03:85, 03:87:02-03:88, 03:91:01-03:93, 03:95-03:98, 03:100-03:101, 03:104-03:118, 03:119:02-03:121N, 03:123-03:127, 03:129-03:134, 03:136-03:142, 03:144-03:152, 03:154-03:166, 03:168-03:177, 03:179-03:189N, 03:191-03:201N, 03:203-03:208N, 03:210-03:219, 03:221-03:239, 03:241-03:261, 03:264-03:266, 03:268-03:270, 03:272-03:275, 03:277N-03:295, 03:297-03:299, 03:301, 03:303-03:329, 03:332-03:334, 03:336-03:337, 03:339-03:343, 04:15:01-04:15:03, 04:17, 04:80, 04:100, 04:224, 04:230, 04:242, 06:02:08, 06:34:02, 08:01:07, 08:02:07, 08:33:02, 12:03:20, 14:05, 14:09, 14:54, 15:77, 16:01:06 *07:30 | B*15:78:03 |
| 73 | 175 bp 205 bp 325 bp | *06:161, 08:06, 16:33 *06:12 *01:90, 01:101-01:102, 01:113, 01:116, 02:06:01-02:06:02, 02:23, 02:36, 02:68, 03:81, 03:175, 03:199, 03:245, 03:317, 04:108, 04:178, 05:132, 06:89, 07:123, 07:173, 07:294, 08:113, 12:15, 12:113, 15:02:01:01-15:02:17, 15:02:19-15:02:21, 15:02:23-15:03, 15:05:01-15:13:01:01, 15:15-15:19, 15:21-15:24, 15:26-15:29, 15:31-15:39, 15:41-15:63, 15:67-15:75, 15:78:01-15:101, 15:103-15:104, 15:106-15:109, 15:111-15:128, 15:130-15:133, 16:20 |  |
| 74 | 95 bp 175 bp | *01:03, 01:24, 01:27, 01:78, 03:58, 04:37, 05:85, 07:364 *03:08, 03:29, 03:31, 03:246, 07:20, 07:96:01-07:96:02, 07:263, 15:02:01:01-15:09, 15:12-15:13:01:01, 15:15, 15:18-15:19, 15:21-15:24, 15:26, 15:28-15:37, 15:39-15:42, 15:44:01-15:62, 15:64-15:106, 15:108-15:112, 15:114-15:123, 15:125-15:133, 16:70 |  |
| 75 | 145 bp 195 bp | *07:57 *02:12, 02:49, 02:55:01-02:55:02, 02:115w, 03:15, 03:27, 03:38:01-03:38:02, 03:69, 03:130, 03:136, 03:163, 03:246w, 03:297, 04:03:01, 04:03:03, 04:06, 04:16, 04:80, 04:107, 04:147, 04:160, 04:171, 04:190, 04:256, 06:70:01-06:70:02, 06:132:01-06:132:02, 07:20, 07:96:01-07:96:02, 07:127, 07:263, 07:390w, 07:562w, 12:140w, 15:03, 15:16, 15:25 |  |
| 78 | 195 bp 235 bp 310 bp 360 bp | *06:12, 07:349 *06:16N *01:02:01:01-01:02:41, 01:05-01:20, 01:22-01:23, 01:25-01:33, 01:35-01:53, 01:56N-01:77, 01:79:01-01:89N, 01:91-01:96, 01:98N-01:101, 01:104-01:114, 01:117N-01:126, 01:128-01:130, 02:19, 02:23, 03:03:01:01-03:11:02, 03:13:01-03:13:02, 03:17:01-03:18:02, 03:20N-03:25, 03:27-03:32, 03:34-03:35:02, 03:37:01-03:39, 03:41:01-03:41:02, 03:44-03:57:02, 03:59, 03:61-03:70, 03:72-03:83, 03:85-03:88, 03:90-03:94, 03:96-03:101, 03:103-03:107, 03:109, 03:111-03:118, 03:120, 03:122-03:131, 03:133-03:138, 03:140-03:145, 03:147-03:168, 03:170-03:174, 03:176-03:189N, 03:191-03:193, 03:195-03:196, 03:202-03:215, 03:217-03:220, 03:223, 03:227-03:237, 03:239, 03:241-03:244Q, 03:246-03:247, 03:249, 03:251-03:257, 03:259-03:263:02, 03:265N-03:266, 03:268-03:270, 03:272-03:278, 03:280-03:297, 03:302-03:306, 03:308-03:314, 03:316N, 03:318N-03:328, 03:331, 03:333-03:334, 03:337, 03:339-03:343, 04:89, 04:135, 04:210, 04:237, 05:47, 05:107, 06:107, 06:179, 07:29:01-07:29:02, 07:177, 07:184, 07:248, 14:54-14:55, 15:06:01-15:06:03, 15:18, 15:40, 15:55, 15:102 *06:15 |  |
| 80 | 70 bp 170 bp 210 bp 235 bp | *03:34, 03:142, 03:261, 03:272 *01:90w, 02:06:01-02:06:02, 02:47, 05:132, 06:168, 12:15, 14:19, 15:08, 15:74 *03:19, 03:44, 03:102, 03:307 *03:88 | A*26:41 B*07:78w, B*13:18w, B*13:31w, B*13:41w, B*13:73w, B*14:19, B*15:73w, B*15:303w, B*54:10w, B*54:20w, B*54:33w, B*55:09w, B*55:21w, B*55:37w, B*55:52w, B*56:43w, B*56:51w B*56:37w |
| 81 | 65 bp 180 bp 440 bp | *07:06, 07:18-07:19, 07:330:02, 07:517, 08:102 *04:30 *01:90, 02:06:01-02:06:02, 02:47, 03:19, 03:102, 03:307, 03:318N, 04:178, 05:132, 06:168, 07:289, 12:15, 15:11, 15:16-15:17, 15:43, | B*46:11, B*46:18, B*56:08, B*56:14 |
| 83 | 235 bp 315 bp 360 bp | *06:16N *02:06:01-02:06:02, 02:47, 03:19, 03:102, 03:307, 04:178, 05:132, 06:168, 07:289, 12:15, 15:02:01:01-15:02:08, 15:02:10-15:03, 15:07-15:08, 15:10:01-15:13:01:01, 15:15-15:18, 15:21, 15:26, 15:28, 15:31-15:35, 15:37-15:39, 15:41-15:45, 15:47-15:53, 15:56-15:58, 15:60, 15:62-15:63, 15:67-15:68, 15:71-15:75, 15:78:01-15:89, 15:91-15:101, 15:103, 15:106, 15:109, 15:112-15:114, 15:118-15:124, 15:127, 15:129-15:133 *06:15 |  |
| 86 | 140 bp 235 bp | *07:72 *02:14:01-02:14:02, 02:107, 03:39, 03:67, 05:43, 06:05, 06:44, 06:124, 06:138, 07:39-07:40, 07:177, 07:210, 07:238, 07:328, 07:335, 07:563, 08:37, 15:23, 15:63, 16:21, 16:80 | B*08:15 |
| 88 | 85 bp 195 bp 545 bp | *04:23, 04:108 *04:31 *01:85, 08:22, 08:56, 08:102, 15:29, 15:87 |  |
| 90 | 205 bp 350 bp | *07:66 *17:01:01:02-17:01:03, 17:01:05-17:01:13, 17:03:01:01-17:03:02, 17:04?, 17:05, 17:06?-17:16?, 17:17, 17:18?-17:29?, 17:30, 17:31?-17:34? |  |
| 91 | 70 bp 105 bp 170 bp 245 bp 270 bp | *17:02, 17:04?, 17:06?-17:16?, 17:18?-17:29?, 17:31?-17:34? *07:67 *07:48 *04:01:43, 05:01:17, 05:09:02, 05:17, 07:13, 07:28, 07:55N, 07:289, 07:536, 08:15:02, 08:72:02 *07:47, 07:461 | B*35:03:15, B*37:04:02 |
| 92 | 155 bp 300 bp 485 bp | *03:135, 04:20 *17:03:01:01-17:03:02, 17:04?, 17:06?-17:16?, 17:17, 17:18?-17:29?, 17:31?-17:34? *07:419 |  |
| 93 | 90 bp 210 bp 290 bp | *17:01:01:02-17:01:13, 17:04?, 17:05, 17:06?-17:16?, 17:18?-17:29?, 17:30, 17:31?-17:34? *03:271, 14:06, 14:15, 14:53, 14:77 *14:11 |  |

4The following HLA-C alleles can be distinguished by the different sizes of the specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| C*01:06, 01:11 | 3 | C*05:10, 05:55 | 21 |
| C*01:15, 01:20, 01:60 | 6 | C*05:14, 05:16, 05:19 | 37 |
| C*02:51, 02:57 | 31 | C*06:08, 06:13 | 43 |
| C*03:09, 03:23, 03:80:02, 03:287 | 18 | C*06:15-06:16N | 78, 83 |
| C*03:21-03:22Q, 03:116:01 | 18 | C*07:22, 07:35, 07:404 | 56 |
| C*03:34, 03:44 | 80 | C*07:25, 07:137:01, 07:138, 07:352 | 56 |
| C*03:37:01-03:37:02, 03:46 | 11 | C*07:33N, 07:51 | 55 |
| C*04:08, 04:09N | 27 | C*07:37, 07:54 | 2 |
| C*04:16, 04:256 | 11 | C*07:47-07:48 | 91 |
| C*04:23, 04:31 | 88 | C*07:140-07:141:02, 07:304, 07:512, 07:514 | 5 |
| C*04:40, 04:223:02 | 31 | C*12:06,12:81 | 65 |
| C*05:06, 05:15 | 36 | C*12:07 ,12:46N | 16 |

‘?’, nucleotide sequence information not available for the primer matching sequence.

‘w’, might be weakly amplified.