Gel Documentation Form and Worksheet

HLA-C low resolution 	Lot No: 2G8	Expiry Date: 2020-12-01

(101.601-24/12, -24u/12u)

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                              Failed lanes: ____________                         Comments:__________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

Primer mix 2 will for most C*02 alleles give rise to two specific PCR fragments.

Primer mix 4 will for most C*03 alleles give rise to two specific PCR fragments.

Primer mix 7 will for most C*05 alleles give rise to two specific PCR fragments.

Primer mix 9 will for most C*07 alleles give rise to two specific PCR fragments.

Primer mix 10 will for most C*08 alleles give rise to multiple specific PCR fragments.

Primer mix 12 will for most C*12 alleles give rise to two specific PCR fragments.

Primer mix 17 will for most C*15 alleles give rise to two specific PCR fragments.

Primer mix 25 will for most C*04 alleles give rise to two specific PCR fragments.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 12, 13, 19 and 30 may have tendencies of unspecific amplifications.

Primer mixes 2, 15, 16, 18, 20, 25, 28 and 30 have a tendency to giving rise to primer oligomer formation.

Primer mixes 3, 6, 10 may give rise to a lower yield of HLA-specific PCR product than the other HLA-C low primer mixes.

Primer mix 14 might faintly amplify most C*01 and the C*14 alleles.

Primer mix 22 might generate a false band of about 500 base pairs. This band should be disregarded when interpreting HLA-C low resolution typings.

Primer mix 32 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-C alleles listed on the IMGT/HLA web page 2017-October-27, release 3.30.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The C*14:06, 14:09, 14:28:01-14:28:02 and 14:63 and the A*29:86, A*30:62, A*33:63, B*35:252 and B*39:114 alleles give rise to identical amplification patterns with the HLA-C low resolution primer set. These alleles are separated by the HLA-A low primer set and respective by the HLA-B low primer set.

‘w’,  might be weakly amplified.

‘?’, nucleotide sequence information not available for the primer matching sequence.