Gel Documentation Form and Worksheet

HLA-C low resolution screening 	Lot No: 2G0	Expiry Date: 2020-09-01

(101.603-24/12,-24u/12u)

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 2 will for most C*02 alleles give rise to two specific PCR fragments.

Primer mix 4 will for most C*03 alleles give rise to two specific PCR fragments.

Primer mix 9 will for most C*07 alleles give rise to two specific PCR fragments.

Primer mix 10 will for most C*08 alleles give rise to multiple specific PCR fragments.

Primer mix 17 will for most C*15 alleles give rise to two specific PCR fragments.

Primer mixes 12 and 23 may have tendencies of unspecific amplifications.

Primer mixes 2, 16 and 18 have a tendency to giving rise to primer oligomer formation.

Primer mixes 3, 6 and 10 may give rise to a lower yield of HLA-specific PCR product than the other HLA-C low Screening primer mixes.

Primer mix 14 might faintly amplify most C*01 and the C*14 alleles.

Primer mix 24 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs.

1HLA-C alleles listed on the IMGT/HLA web page 2017-October-27, release 3.30.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following alleles give rise to identical amplification patterns with the HLA-C low screening resolution primer set. These alleles can be separated by the respective high resolution primer sets.

| Alleles | Alleles |
| --- | --- |
| C*12:02:14-12:02:15, 12:18:02, B*67:02 | C*14:06, 14:08-14:09, 14:28:01-14:28:02, 14:53, 14:63, A*29:86, A*30:62, A*33:63, B*35:252, B*39:114 |
| C*12:49, 12:123, B*57:01:24 |  |

‘w’, might be weakly amplified.

‘?’, nucleotide sequence of the primer matching region not known.