# HLA-C low resolution screening 	Lot No: 4N2	Expiry Date: 2025-07-01

# (101.603-24/12,-24u/12u)

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

# Interpretation:___________                            Failed lanes: ____________                              Comments:________________

Gel Picture

## PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mix 2 will give rise to two specific PCR fragments for most C*02 alleles.

Primer mix 4 will give rise to two specific PCR fragments for most C*03 alleles.

Primer mix 9 will give rise to two specific PCR fragments for most C*07 alleles.

Primer mix 10 will give rise to multiple specific PCR fragments for most C*08 alleles.

Primer mix 17 will give rise to two specific PCR fragments for most C*15 alleles.

Primer mixes 12 and 23 may have tendencies of unspecific amplifications.

Primer mixes 2, 16 and 18 have a tendency to giving rise to primer oligomer formation.

Primer mixes 6, 10 and 16 may give rise to a lower yield of HLA-specific PCR product than the other HLA-C low Screening primer mixes.

Primer mix 14 might faintly amplify most C*01 and the C*14 alleles.

In primer mix 22 the positive control band may be weaker than for other HLA-C low resolution primer mixes.

Primer mix 24 contains a negative control, which will amplify the majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR products generated by the HGH positive control primer pair is 200 base pairs.

1HLA-C alleles listed on the IMGT/HLA web page 2021-April-20, release 3.44.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following alleles give rise to identical amplification patterns with the HLA-C low screening resolution primer set. These alleles can be separated by the respective high resolution primer sets.

| Alleles | Alleles |
| --- | --- |
| C*01:125, 14:58 | C*12:02:29, 12:49, 12:123, B*18:01:42, B*57:01:24 |
| C*01:169:02, 03:94 | C*12:03:34:01-12:03:34:02, B*15:510, B*35:475 |
| C*03:16, 03:536, B*46:77 | C*14:06, 14:08-14:09, 14:28:01-14:28:02, 14:63, 14:105Q, A*29:86, A*30:62, A*31:187, A*33:63, B*08:249, B*18:64, B*35:183, 35:252, 35:502, B*39:114, 39:140, B*40:138, B*41:59 |
| C*05:206, 17:05 | C*14:110, B*07:239, B*14:03 |
| C*12:02:14-12:02:15, 12:02:19, 12:02:30, 12:02:32, 12:02:34, 12:02:37, 12:02:40, 12:18:02, 12:222, 12:224, B*67:02:01:01-67:02:01:02 |

Abbreviations

w: might be weakly amplified.

?: nucleotide sequence of the primer matching region not known.