# Gel Documentation Form and Worksheet

# HLA-C*01 (101.621-12/12u) 	Lot No: 4G0	Expiry Date: 2022-05-01

# Name:__________________

# Sample ID:______________

# DNA Conc.(ng/ul):_________

# Test  Date:_______________

# Tested By:_______________

# Review Date:_______________

# Reviewed By:_______________

# Interpretation:_______________	                    Failed lanes: ____________                           Comments:________________

Gel Picture

## PHOTO DOCUMENT

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 6, 7, 14, 18 and 22 may have tendencies of unspecific amplifications.

Primer mixes 1 and 13 may give rise to a lower yield of HLA-specific PCR product than the other HLA-C*01 primer mixes.

Primer mix 25 has a tendency to giving rise to primer oligomer formation.

Primer mix 32 contains a negative control, which will amplify more than 95% of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 430 base pairs

1HLA-C*01 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.25.0, July 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-C*01 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-C*01 alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- |
| 2 | 90 bp 270 bp | *01:03, 01:24, 01:78, 01:146 *01:15:01 | *03:58, 04:37, 05:85, 07:364 |
| 5 | 105 bp 150 bp 200 bp | *01:38 *01:20 *01:06 |  |
| 6 | 195 bp 230 bp | *01:07:01-01:07:02 *01:37N, 01:83 | *06:43:01, 14:24:02 *14:35N |
| 7 | 70 bp 150 bp 195 bp | *01:67 *01:20 *01:08 | *03:03:10, 03:04:28, 04:01:11, 06:02:21, 07:01:58, 07:02:36, 12:03:36, 16:01:19, A*01:01:33, A*02:01:29, A*03:01:42, A*11:01:40, A*24:07:02, A*26:01:09, A*32:01:09, A*33:01:07, A*68:01:06, B*07:02:21, B*13:02:03, B*15:01:39, B*27:05:06, B*35:08:07, B*40:01:10, B*40:02:11, B*44:02:37, B*44:03:08, B*51:01:24, B*73:01-73:02, B*82:02:02 |
| 8 | 210 bp 260 bp | *01:04, 01:09, 01:22, 01:35 *01:143N | *03:302, 06:23, 06:179, 07:177, 12:178, 15:37, B*40:243, B*54:38 |
| 9 | 160 bp 225 bp | *01:52 *01:10, 01:83 | B*40:243 |
| 10 | 210 bp 255 bp 290 bp | *01:22, 01:35 *01:30 *01:11 | *03:302, 07:177, 15:37, B*40:243, B*54:38 |
| 11 | 140 bp 355 bp | *01:39 *01:12:01-01:12:02, 01:34, 01:79:01-01:79:02, 01:97, 01:101-01:102, 01:114, 01:131, 01:137N |  |
| 12 | 80 bp 155 bp 255 bp | *01:84 *01:13 *01:82 | *03:213 *02:51, 03:87:01-03:87:02, 04:223:01-04:223:02, 05:09:01-05:09:03, 05:17, 05:52, 07:130, 08:15:01-08:15:02, 08:51, 12:144, 12:185, 16:27, B*15:33, B*15:248 |
| 14 | 120 bp 240 bp | *01:41 *01:17, 01:21, 01:23, 01:69N, 01:128 | *07:316, 07:338, 07:579, 12:215 |
| 15 | 115 bp 230 bp | *01:42, 01:73 *01:16, 01:18, 01:74 | B*13:102, B*15:393, B*46:60, B*51:129 |
| 16 | 130 bp 255 bp 295 bp | *01:43 *01:19 *01:23, 01:58 | *07:316, 07:338, 07:579, A*01:24 |
| 17 | 75 bp 255 bp | *01:24-01:25 *01:82 |  |
| 18 | 195 bp 260 bp 295 bp 345 bp | *01:26 *01:34, 01:146 *01:145N *01:36, 01:49:01, 01:55 |  |
| 19 | 100 bp 265 bp | *01:27 *01:30, 01:45 |  |
| 20 | 80 bp 110 bp 285 bp | *01:84 *01:28 *01:56N | *03:213 *03:59, 03:123, 06:157 |
| 21 | 125 bp 160 bp 245 bp | *01:33 *01:80 *01:29, 01:69N |  |
| 22 | 110 bp 250 bp 335 bp 355 bp | *01:40 *01:32:01-01:32:02 *01:50, 01:131 *01:05, 01:36, 01:55, 01:79:01-01:79:02, 01:120 | *06:110 |
| 23 | 90 bp 120 bp 235 bp | *01:04, 01:54, 01:103 *01:44 *01:31, 01:35, 01:107 | *06:23, 12:178, 14:45, 16:18 *03:302 |
| 24 | 90 bp 165 bp 260 bp | *01:86N, 01:103 *01:66 *01:143N |  |
| 25 | 85 bp 240 bp 270 bp | *01:99 *01:16 *01:70 | A*24:112, B*13:102, B*51:129 |
| 26 | 155 bp 230 bp | *01:117N *01:74, 01:98N |  |
| 27 | 350 bp 545 bp | *01:14, 01:59, 01:118 *01:85 | *04:37, 05:85, 05:107, 06:23, 06:179, 06:204, 15:37, 15:102 *04:277, 08:22, 08:56, 08:102, 08:154, 15:29, 15:87 |
| 28 | 155 bp 295 bp 325 bp 360 bp | *01:35, 01:107, 01:131 *01:145N *01:81 *01:49:01-01:50, 01:131 |  |
| 29 | 135 bp 295 bp 350 bp | *01:109N *01:89N *01:14, 01:59, 01:118 | *04:37, 05:85, 05:107, 06:23, 06:179, 06:204, 15:37, 15:102 |
| 30 | 125 bp 235 bp | *01:93 *01:121Q |  |
| 31 | 155 bp 265 bp | *01:117N *01:14 | *15:104 |

4The HLA-C*01 primer set cannot separate the C*01:21 from the C*12:215 allele. These alleles can be distinguished by the HLA-C low resolution kit and/or by the HLA-C*12 high resolution kit.

The HLA-C*01 primer set cannot separate the C*01:123 and the C*01:125 from the C*03:86, 03:94 and 03:99:01-03:99:02 alleles. These alleles can be distinguished by the HLA-C low resolution kit and/or by the HLA-C*03 high resolution kit.

5The following HLA-C*01 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| C*01:06, 01:38 | 5 | C*01:29, 01:33 | 21 |
| C*01:10, 01:52 | 9 | C*01:31, 01:44 | 23 |
| C*01:17, 01:41, 01:128 | 14 | C*01:32:01-01:32:02, 01:40 | 22 |
| C*01:18, 01:42 | 15 | C*01:39, 01:137N | 11 |
| C*01:19, 01:43, 01:58 | 16 | C*01:70, 01:99 | 25 |
| C*01:27, 01:45 | 19 | C*01:89N, 01:109N | 29 |
| C*01:28, 01:56N | 20 | C*01:93, 01:121Q | 30 |

Changes in revision R01 compared to R00:

1.	The expiration date has been altered due to extension of shelf-life.