Gel Documentation Form and Worksheet

HLA-C*12 (101.624-12/12u) 	Lot No:  6L7	Expiry Date: 2024-10-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                           Failed lanes: ____________                          Comments:________________

Gel Picture

PHOTO DOCUMENT

Abbrevations

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 5 and 26 may give rise to a lower yield of HLA-specific PCR product than the other C*12 primer mixes

Primer mixes 24 and 28 have tendencies of giving rise to primer oligomer formation.

Primer mixes 30 and 40 may have tendencies of unspecific amplifications.

Primer mix 48 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1HLA-C*12 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.27.0, January 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-C*12 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-C*12 alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- |
| 6 | 75 bp 150 bp 415 bp | *12:48, 12:102 *12:06 *12:08, 12:81, 12:188, 12:287 | *01:118, 01:128, 01:195, 03:08, 03:29, 03:31, 03:246, 04:112, 04:169, 05:36, 06:44, 06:252, 14:73, 16:64, 16:70, 16:87 |
| 8 | 95 bp 155 bp 195 bp 245 bp | *12:15 *12:40 *12:80N *12:07 | *15:02:14 *16:14 B*35:310 |
| 13 | 105 bp 150 bp | *12:31:01-12:31:02 *12:10:01-12:10:02, 12:155Q-12:156 | *04:01:05, 04:01:75, 04:01:124, 05:106:02, 08:01:19, 08:160 *04:01:05, 04:01:75, 04:01:124, 14:02:08, 14:02:21, 14:03:05 |
| 14 | 100 bp 150 bp | *12:16:01, 12:147, 12:227, 12:279 *12:11, 12:60, 12:118 | *01:02:34, 01:21, 02:42, 02:107, 02:152, 04:140, 04:166:01, 04:220, 05:98, 05:197, 06:02:72, 06:05w, 07:01:74, 07:02:09, 08:14, 08:80, 08:103, 15:63, 15:113, 16:80, B*15:436, B*67:02:01:01-67:02:01:02 |
| 16 | 185 bp 225 bp | *12:13:01:01-12:13:01:02 *12:14:01-12:14:02, 12:176 | *01:60, 02:180, 04:58, 04:160, 04:368, 05:23, 05:62, 05:134, 05:143, 05:151, 06:118, 08:07, 08:47, 08:104, 08:188, 14:17, 15:65, 17:01:01:02-17:31, 17:33-17:34, 17:36-17:54 |
| 17 | 130 bp 565 bp | *12:99:02, 12:159 *12:03:04, 12:03:09, 12:195:01 | B*40:02:21w *01:02:18, 03:03:40, 06:02:38, 07:447, 14:02:08, B*27:05:27, B*57:01:24 |
| 18 | 145 bp 175 bp 245 bp 270 bp | *12:17, 12:27 *12:35, 12:201 *12:17, 12:27 *12:35 | *04:12 *03:53, A*68:226 A*68:166, B*07:297, B*38:60, B*55:92 |
| 20 | 105 bp 175 bp 230 bp | *12:46N *12:22, 12:58, 12:94, 12:252 *12:19, 12:139Q, 12:158 | *01:129, 04:52, 04:55, 04:405, 05:55, 14:10, 14:48, 15:12, 15:144 *01:31, 14:38 |
| 22 | 100 bp 135 bp 590 bp | *12:15, 12:23 *12:99:01-12:99:02, 12:235 *12:21, 12:203 | *15:02:14 *03:03:40, B*40:02:21w *05:106:02, 08:01:19 |
| 23 | 105 bp 140 bp 185 bp | *12:205 *12:26, 12:63 *12:172, 12:201 | *05:235, 06:56, 07:620, 15:147, A*02:362, A*26:85, A*66:41, B*13:80, B*18:116, B*35:326, B*44:38 *03:465, 07:470, 16:36 *03:477, 03:496, 04:337 |
| 24 | 135 bp 185 bp 425 bp | *12:99:01, 12:235 *12:43 *12:28, 12:135 | *03:03:40 *04:01:05, 04:01:124, 06:02:38, 06:76:02, 07:447 |
| 25 | 80 bp 155 bp 430 bp | *12:39N *12:02:06, 12:02:08, 12:02:12, 12:21, 12:118, 12:149 *12:167, 12:243 | *02:02:44, 04:01:05, 04:01:124, 05:106:02, 07:413, 07:422, 08:01:19, 08:02:02, B*27:05:27, B*40:02:21 *02:12, 02:49w, 02:55:01w-02:55:02w, 02:115, 04:226w, 15:03w, 15:16w, 15:25 |
| 27 | 100 bp 150 bp 175 bp 215 bp 295 bp | *12:30 *12:03:19, 12:03:32, 12:155Q *12:94 *12:36, 12:153 *12:101 | *07:214, 07:429 *01:04, 14:02:21, 14:03:05, 16:02:13 *01:129, 14:48, A*02:605Q *16:103 *14:84, 16:81, 16:143 |
| 28 | 165 bp 275 bp 350 bp | *12:232N *12:50 *12:45, 12:166 | *01:32:01-01:32:02, 02:56, 03:102, 03:263:01-03:263:02, 03:514, 04:180:01, 05:217, 06:20, 07:81, 07:168, 07:450, 08:123, 08:139, 14:82, 14:92, 15:126, 16:98, 16:102, 16:110 *02:159, 02:161, 05:81, 06:87, 07:24, 07:218, 14:65, 16:13, 16:61 |
| 29 | 125 bp 185 bp 210 bp | *12:38, 12:104N, 12:219N *12:42Q, 12:172 *12:29, 12:86 | *07:820N, 14:93N, B*15:181N, B*57:50 *03:477, 03:496, 04:337, 07:513Q, B*46:51Q *16:119 |
| 30 | 90 bp 200 bp 230 bp | *12:32, 12:102, 12:144, 12:185 *12:162 *12:34 | *02:51, 05:08, 05:52, 05:89, 06:41, 08:29, 08:31, B*15:33, B*15:248 *02:64 |
| 31 | 135 bp 180 bp 240 bp | *12:47, 12:84N, 12:123 *12:42Q, 12:80N *12:164 | *02:133, A*11:197, A*26:67, A*68:95 *07:513Q, B*46:51Q *15:67 |
| 32 | 55 bp 115 bp 180 bp | *12:54, 12:188 *12:37 *12:62 | *01:59, 01:118, 02:65, 03:130, 03:140:01:01-03:140:01:02, 03:243, 04:114, 04:383, 05:20, 06:82, 06:210, 07:49, 07:210, 07:238, 07:247, 07:403, 14:04, 14:64, 14:77, 15:85, 15:181, 16:57, A*03:267, A*68:46, B*07:253 *07:204:01, 07:482, A*02:211:01, A*02:594, A*02:817, A*24:261, A*24:445N, A*68:76:01-68:76:02 A*02:335 |
| 33 | 135 bp 210 bp | *12:105N, 12:219N *12:100 | *07:820N |
| 34 | 85 bp 235 bp 260 bp | *12:171 *12:109 *12:125 | *02:170, 03:171, 03:211:01, 04:144, 05:93, 06:73, 08:20, 08:40, 15:221 |
| 35 | 195 bp 260 bp | *12:110, 12:143, 12:278 *12:125 | *01:173, 03:411, 04:375, 04:380, 06:227, 15:154 |
| 37 | 90 bp 225 bp 285 bp | *12:163 *12:108 *12:73 | B*14:51, B*53:34, B*58:21 |
| 38 | 105 bp 145 bp | *12:15, 12:113, 12:282, 12:309 *12:156 | *15:02:14, B*40:02:21, B*57:01:24 *04:01:05, 04:01:124, 14:02:08 |
| 39 | 200 bp 225 bp | *12:136 *12:59, 12:82 | *04:280w *07:102, 07:351, A*25:55, B*07:13, B*07:15, B*07:160, B*42:18, B*67:02:01:01-67:02:01:02 |
| 41 | 215 bp 240 bp | *12:90, 12:148N *12:164 | B*56:08 *15:67 |
| 44 | 145 bp 290 bp | *12:16:01, 12:147, 12:195:02, 12:217 *12:160 | *02:14:01-02:14:02, 02:107, 02:164, 04:42:01-04:42:02, 04:220, 05:43, 06:02:72, 06:05, 07:01:74, 07:02:09, 08:37, 15:23:01-15:23:02, 15:63, 15:138, 15:158, 16:21, 16:80 |
| 46 | 175 bp 190 bp | *12:165, *12:143, 12:278 | *02:148 *01:173, 03:411, 04:375, 04:380, 06:227, 15:154 |

4The following alleles give rise to identical amplification patterns with the HLA-C*12 primer set. These alleles can be distinguished by the HLA-C low resolution kit and the HLA-C*05, HLA-C*07, HLA-C*08 or HLA-C*16, kit, respectively.

Alleles

C*12:09, C*05:16, C*05:85, C*05:107, C*16:88

C*12:33, C*16:91

C*12:181, C*07:723

C*12:208, C*12:222, C*12:233, C*12:251, C*12:258, C*08:05, C*08:25, C*16:15:01

5The following HLA-C*12 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| C*12:03:09, 12:159 | 17 | C*12:32, 12:34 | 30 |
| C*12:06, 12:48, 12:81 | 6 | C*12:39N, 12:167 | 25 |
| C*12:11, 12:227 | 14 | C*12:45, 12:50, 12:232N | 28 |
| C*12:23, 12:203 | 22 | C*12:46N, 12:139Q | 20 |
| C*12:29, 12:38 | 29 | C*12:86, 12:104N | 29 |
| C*12:30, 12:36 | 27 | C*12:109, 12:171 | 34 |