Gel Documentation Form and Worksheet

HLA-C*12 (101.624-12/12u) 	Lot No:  7K1	Expiry Date: 2024-02-01

Name:__________________

Sample ID:______________

DNA Conc.(ng/ul):_________

Test  Date:_______________

Tested By:_______________

Review Date:_______________

Reviewed By:_______________

Interpretation:___________                           Failed lanes: ____________                          Comments:________________

Gel Picture

PHOTO DOCUMENT

Abbrevations

‘ICB’ Internal Control Band,

‘AmpS’ Amplicon Size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

Primer mixes 5 and 26 may give rise to a lower yield of HLA-specific PCR product than the other C*12 primer mixes

Primer mix 28 has a most pronounced tendency to giving rise to primer oligomer formation.

Primer mixes 30 and 40 may have tendencies of unspecific amplifications.

Primer mix 48 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1HLA-C*12 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.27.0, January 2017.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-C*12 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-C*12 alleles | Other amplified HLA Class I alleles |
| --- | --- | --- | --- |
| 6 | 75 bp 150 bp 415 bp | *12:48, 12:102 *12:06 *12:08, 12:81, 12:188, 12:287 | *01:118, 01:128, 03:08, 03:29, 03:31, 03:246, 04:112, 04:169, 05:36, 06:44, 06:252, 14:73, 16:64, 16:70, 16:87 |
| 8 | 95 bp 155 bp 195 bp 245 bp | *12:15 *12:40 *12:80N *12:07 | *15:02:14 *16:14 B*35:310 |
| 13 | 105 bp 150 bp | *12:31 *12:10:01-12:10:02, 12:155-12:156 | *04:01:05, 04:01:75, 05:106:02, 08:01:19, 08:160 *04:01:05, 04:01:75, 14:02:08, 14:02:21 |
| 14 | 100 bp 150 bp | *12:16:01, 12:147, 12:227, 12:279 *12:11, 12:60, 12:118 | *01:02:34, 01:21, 02:42, 02:107, 02:152, 04:140, 04:166:01, 04:220, 05:98, 05:197, 06:02:72, 06:05w, 07:01:74, 07:02:09, 08:14, 08:80, 08:103, 15:63, 15:113, 16:80, B*15:436, B*67:02:01:01-67:02:01:02 |
| 16 | 185 bp 225 bp | *12:13:01:01-12:13:01:02 *12:14:01-12:14:02, 12:176 | *01:60, 02:180, 04:58, 04:160, 04:368, 05:23, 05:62, 05:134, 05:143, 05:151, 06:118, 08:07, 08:47, 08:104, 08:188, 14:17, 15:65, 17:01:01:02-17:31, 17:33-17:34, 17:36-17:51:01:02 |
| 17 | 130 bp 565 bp | *12:99:02, 12:159 *12:03:04, 12:03:09, 12:195:01 | B*40:02:21w *01:02:18, 03:03:40, 06:02:38, 07:447, 14:02:08, B*27:05:27, B*57:01:24 |
| 18 | 145 bp 175 bp 245 bp 270 bp | *12:17, 12:27 *12:35, 12:201 *12:17, 12:27 *12:35 | *04:12 *03:53, A*68:226 A*68:166, B*07:297, B*38:60, B*55:92 |
| 20 | 105 bp 175 bp 230 bp | *12:46N *12:22, 12:58, 12:94, 12:252 *12:19, 12:139, 12:158 | *01:129, 04:52, 04:55, 04:405, 05:55, 14:10, 14:48, 15:12, 15:144 *01:31, 14:38 |
| 22 | 100 bp 135 bp 590 bp | *12:15, 12:23 *12:99:01-12:99:02, 12:235 *12:21, 12:203 | *15:02:14 *03:03:40, B*40:02:21w *05:106:02, 08:01:19 |
| 23 | 105 bp 140 bp 185 bp | *12:205 *12:26, 12:63 *12:172, 12:201 | *06:56, 07:620, 15:147, A*02:362, A*26:85, B*13:80, B*18:116, B*35:326, B*44:38 *03:465, 07:470, 16:36 *03:477, 03:496, 04:337 |
| 24 | 135 bp 185 bp 425 bp | *12:99:01, 12:235 *12:43 *12:28, 12:135 | *03:03:40 *04:01:05, 06:02:38, 06:76:02, 07:447 |
| 25 | 80 bp 155 bp 430 bp | *12:39N *12:02:06, 12:02:08, 12:02:12, 12:21, 12:118, 12:149 *12:167, 12:243 | *02:02:44, 04:01:05, 05:106:02, 07:413, 07:422, 08:01:19, 08:02:02, B*27:05:27, B*40:02:21 *02:12, 02:49w, 02:55:01w-02:55:02w, 02:115, 04:226w, 15:03w, 15:16w, 15:25 |
| 27 | 100 bp 150 bp 175 bp 215 bp 295 bp | *12:30 *12:03:19, 12:03:32, 12:155Q *12:94 *12:36, 12:153 *12:101 | *07:214, 07:429 *01:04, 14:02:21, 16:02:13 *01:129, 14:48, A*02:605Q *16:103 *14:84, 16:81, 16:143 |
| 28 | 165 bp 275 bp 350 bp | *12:232N *12:50 *12:45, 12:166 | *01:32:01-01:32:02, 02:56, 03:102, 03:263:01-03:263:02, 04:180:01, 05:217, 06:20, 07:81, 07:168, 07:450, 08:123, 08:139, 14:82, 14:92, 15:126, 16:98, 16:102, 16:110 *02:159, 02:161, 05:81, 06:87, 07:24, 07:218, 14:65, 16:13, 16:61 |
| 29 | 125 bp 185 bp 210 bp | *12:38, 12:104N, 12:219N *12:42Q, 12:172 *12:29, 12:86, 16:119 | *07:820N , 14:93N, B*15:181N, B*57:50 *03:477, 03:496, 04:337, 07:513Q, B*46:51Q |
| 30 | 90 bp 200 bp 230 bp | *12:32, 12:102, 12:144, 12:185 *12:162 *12:34 | *02:51, 05:08, 05:52, 05:89, 06:41, 08:29, 08:31, B*15:33, B*15:248 *02:64 |
| 31 | 135 bp 180 bp 240 bp | *12:47, 12:84N, 12:123 *12:42Q, 12:80N *12:164 | *02:133, A*11:197, A*26:67, A*68:95 *07:513Q, B*46:51Q *15:67 |
| 32 | 55 bp 115 bp 180 bp | *12:54, 12:188 *12:37 *12:62 | *01:59, 01:118, 02:65, 03:130, 03:140, 03:243, 04:114, 04:383, 05:20, 06:82, 06:210, 07:49, 07:210, 07:238, 07:247, 07:403, 14:04, 14:64, 14:77, 15:85, 15:181, 16:57, A*03:267, A*68:46, B*07:253 *07:204:01, 07:482, A*02:211:01, A*02:594, A*02:817, A*24:261, A*24:445N, A*68:76:01-68:76:02 A*02:335 |
| 33 | 135 bp 210 bp | *12:105N, 12:219N *12:100 | *07:820N |
| 34 | 85 bp 235 bp 260 bp | *12:171 *12:109 *12:125 | *02:170, 03:171, 03:211:01, 04:144, 05:93, 06:73, 08:20, 08:40 |
| 35 | 195 bp 260 bp | *12:110, 12:143, 12:278 *12:125 | *01:173, 03:411, 04:375, 04:380, 06:227, 15:154 |
| 37 | 90 bp 225 bp 285 bp | *12:163 *12:108 *12:73 | B*14:51, B*53:34, B*58:21 |
| 38 | 105 bp 145 bp | *12:15, 12:113, 12:282 *12:156 | *15:02:14, B*40:02:21, B*57:01:24 *04:01:05, 14:02:08 |
| 39 | 200 bp 225 bp | *12:136 *12:59, 12:82 | *04:280w *07:102, 07:351, A*25:55, B*07:13, B*07:15, B*07:160, B*42:18, B*67:02:01:01-67:02:01:02 |
| 41 | 215 bp 240 bp | *12:90, 12:148N *12:164 | B*56:08 *15:67 |
| 44 | 145 bp 290 bp | *12:16:01, 12:147, 12:195:02, 12:217 *12:160 | *02:14:01-02:14:02, 02:107, 02:164, 04:42:01-04:42:02, 04:220, 05:43, 06:02:72, 06:05, 07:01:74, 07:02:09, 08:37, 15:23:01-15:23:02, 15:63, 15:138, 15:158, 16:21, 16:80 |
| 46 | 175 bp 190 bp | *12:165, *12:143, 12:278 | *02:148 *01:173, 03:411, 04:375, 04:380, 06:227, 15:154 |

4The following alleles give rise to identical amplification patterns with the HLA-C*12 primer set. These alleles can be distinguished by the HLA-C low resolution kit and the HLA-C*05, HLA-C*07, HLA-C*08 or HLA-C*16, kit, respectively.

Alleles

C*12:09, C*05:16, C*05:85, C*05:107, C*16:88

C*12:33, C*16:91

C*12:181, C*07:723

C*12:208, C*12:222, C*12:233, C*12:251, C*12:258, C*08:05, C*08:25, C*16:15:01

5The following HLA-C*12 alleles can be distinguished by the different sizes of the HLA-specific PCR product:

| Alleles | Primer mix | Alleles | Primer mix |
| --- | --- | --- | --- |
| C*12:03:09, 12:159 | 17 | C*12:32, 12:34 | 30 |
| C*12:06, 12:48, 12:81 | 6 | C*12:39N, 12:167 | 25 |
| C*12:11, 12:227 | 14 | C*12:45, 12:50, 12:232N | 28 |
| C*12:23, 12:203 | 22 | C*12:46N, 12:139 | 20 |
| C*12:29, 12:38 | 29 | C*12:86, 12:104N | 29 |
| C*12:30, 12:36 | 27 | C*12:109, 12:171 | 34 |