# Gel Documentation Form and Worksheet

# HLA-C*17 (101.628-06/06u) 	Lot No: 6K3	Expiry Date: 2023-12-01

# Name:___________________

# Sample ID:_______________

# DNA Conc.(ng/ul):_________

# Test Date: ______________

# Tested By: ______________

# Review Date:________________

# Reviewed By:________________

# Interpretation:___________                          Failed lanes: ____________                         Comments:________________

Gel Picture

## PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon size

Notes:

Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

HLA-specific PCR products shorter than 125 base pairs have a lower intensity and are less sharp than longer PCR products.

In order to have as little co-amplification as possible of non-HLA-C*17 alleles by the primer pairs of the HLA-C*17 subtyping kit, many of the specific PCR products are shorter than 100 base pairs.

Primer mix 15 may give rise to a lower yield of HLA-specific PCR product than the other HLA-C*17 primer mixes in the C*07:02 alleles.

Primer mix 16 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

1HLA-C*17 alleles in bold lettering are listed as confirmed alleles on the IMGT/HLA web page www.ebi.ac.uk/imgt/hla, release 3.24.0, April 2016.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

3The following HLA-C*17 primer mixes have two or more product sizes:

| Primer Mix | Size of spec. PCR product | Amplified HLA-C*17 alleles | Other amplified HLA-C alleles |
| --- | --- | --- | --- |
| 5 | 155 bp 200 bp | *17:04 *17:15 | *08:184, 15:135 |
| 7 | 125 bp 200 bp | *17:06 *17:15 |  |
| 9 | 85 bp 135 bp 495 bp | *17:27N *17:08 *17:27N |  |
| 10 | 90 bp 495 bp | *17:09, 17:27N *17:27N | B*40:260 |
| 14 | 80 bp 260 bp | *17:17 *17:13 | *05:206, 08:185 |

Abbreviations

w: might be weakly amplified.

?: nucleotide sequence information not available for the primer matching sequence.