PAGE  
September 2009
HLA-Cw*04 
Page 1 of 3
Rev. No: 02
56E
Worksheet can be downloaded from

Worksheet
www.olerup.com









	PHOTO DOCUMENT


Olerup SSP®
Olerup SSP® HLA-Cw*04
Prod. No:
101.612-12 – including Taq polymerase

101.612-12u – without Taq polymerase
Lot No:    
56E
Expiry Date:    
2010-May-01
Name:________________Sample ID:_______________

DNA Extract Date:_________ Conc.(ng/ul):_________

Test  Date: _________Review Date:_______________

Tested By: _________Reviewed By:_______________

Interpretation:_________________________________

Specificities and sizes of the PCR products of the 24 primer mixes used for HLA-Cw*04 SSP typing.
	Primer Mix
	Size of spec. PCR product1
	Size of control band2
	Amplified HLA-Cw*04 alleles
	Other amplified HLA Class I alleles3

	  1
	250 bp
	  800 bp
	04010101-040105, 0403, 0405, 0407-0412, 041501-0418, 0419weakly, 0420, 0423-0432
	010201-0120, 020201-020203, 020205-020206, 0204-0215, 0217, 0219-0221, 050101-0510, 0512-0516, 0518, 0519, 06020101-0603, 0605-0617, 0802, 0805, 0807, 0812, 0815, 120201-1213, 1215-1217, 1221, 140201-1405, 1407N-1409, 1508, 160401, 1801-1803

	  2
	220 bp
	1070 bp
	04010101-040105, 040401-0405, 0407-0409N, 0412weakly, 0413-041502, 0417-0420, 0423-0432
	

	  34
	145 bp
	1070 bp
	04010101-040105, 040401-0405, 0408-041502, 0417-0420, 0423-0426, 0428-0432
	

	  45
	210 bp
	1070 bp
	0403, 0406
	0212

	  5
	250 bp
	1070 bp
	040401-040402faintly, 0406faintly, 0413faintly
	0203, 0216, 0218, 0511, 0517, 0604, 080101-080102, 0803, 0804, 0806, 0808-0811, 0813, 0814, 0816, 1214, 1218, 1220, 1406, 150201-1507, 1509-1521, 1701-1704,

B*5802

	  64
	  95 bp
	1070 bp
	0405
	

	  74
	145 bp
	1070 bp
	0407, 0427, 0432
	0334, 1803

	  8
	270 bp
	1070 bp
	0408weakly
	1515,

B*5802

	
	
	
	
	

	  96,7,8
	180, 220 bp
	1070 bp
	0409N, 0430
	

	109
	195, 220 bp
	1070 bp
	0410, 0411, 0431
	

	11
	180 bp
	1070 bp
	04010101-040105, 0403-0410, 0412-0420, 0423-0426, 0428-0432
	0605

	124,10
	125, 180 bp
	1070 bp
	0411, 0429, 0430
	

	13
	225 bp
	1070 bp
	0412
	

	1411
	155, 190 bp
	1070 bp
	0413, 0416, 0418
	0708

	156,12
	170, 190 bp
	1070 bp
	0414, 0428
	

	164
	130 bp
	1070 bp
	041501
	0305, 0325, 0327, 1409

	
	
	
	
	

	174,6
	  90 bp
	  800 bp
	0417
	

	18
	225 bp
	1070 bp
	0419
	

	19
	155 bp
	1070 bp
	0420
	

	20
	540 bp
	1070 bp
	041502, 0417
	

	214
	  85 bp
	1070 bp
	0423
	

	224
	130 bp
	1070 bp
	0424
	

	234
	  85 bp
	1070 bp
	0425
	

	24
	170 bp
	1070 bp
	0426
	


1Alleles are assigned by the presence of specific PCR product(s). However, the sizes of the specific PCR products may be helpful in the interpretation of HLA-Cw*04 SSP typings.

When the primers in a primer mix can give rise to specific PCR products of more than one length this is indicated if the size difference is 20 base pairs or more. Size differences shorter than 20 base pairs are not given. For high resolution SSP kits the respective length of the specific PCR product(s) of the alleles amplified by these primer mixes are given. 

Nonspecific amplifications, i.e. a ladder or a smear of bands, may sometimes be seen. GC-rich primers have a higher tendency of giving rise to nonspecific amplifications than other primers.

PCR fragments longer than the control bands may sometimes be observed. Such bands should be disregarded and do not influence the interpretation of the SSP typings.

PCR fragments migrating faster than the control bands, but slower than a 400 bp fragment may  be seen in some gel read-outs. Such bands can be disregarded and do not influence the interpretation of the SSP typings.

Some primers may give rise to primer oligomer artifacts. Sometimes this phenomenon is an inherit feature of the primer pair(s) of a primer mix. More often it is due to other factors such as too low amount of DNA in the PCR reactions, taking too long time in setting up the PCR reactions, working at elevated room temperature or using thermal cyclers that are not pre-heated.

2The internal positive control primer pairs amplify segments of the human growth hormone gene. The two different control primer pairs give rise to either an internal positive control band of 1070 base pairs, for most wells, or a band of 800 base pairs, for some wells.

Well number 1 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to help in the correct orientation of the HLA-Cw*04 subtyping.

In addition, well number 17 contains the primer pair giving rise to the shorter, 800 bp, internal positive control band in order to allow kit identification.

In the presence of a specific amplification the intensity of the control band often decreases.

3Due to the sharing of sequence motifs between HLA-Cw alleles some non-HLA-Cw*04 alleles will be amplified by primer mixes 1, 4, 5, 7, 8, 11, 14 and 16.

4Short specific PCR fragments are less intense and not as sharp as longer specific bands.

5Primer mix 4 may give rise to primer dimer formation.

6The primers in primer mixes 9, 15 and 17 have a tendency of giving rise to nonspecific amplifications.

7The specific PCR product generated by the primers in primer mix 9 may be less intense than the other Cw*04 specific PCR products.

8Primer mix 9: Specific PCR fragment of 180 bp in the Cw*0430 allele. Specific PCR fragment of 220 bp in the Cw*0409N allele.

9Primer mix 10: Specific PCR fragmentof 195 bp in the Cw*0431 allele. Specific PCR fragment of 220 bp in the Cw*0410 and Cw*0411 alleles.

10Primer mix 12: Specific PCR fragment of 125 bp in the Cw*0411 and Cw*0429 alleles. Specific PCR fragment of 180 bp in the Cw*0430 allele.

11Primer mix 14: Specific PCR fragment of 155 bp in the Cw*0416 allele. Specific PCR fragment of 190 bp in the Cw*0413, Cw*0418 and Cw*0708 alleles.

12Primer mix 15: Specific PCR fragment of 170 bp in the Cw*0414 allele. Specific PCR fragment of 190 bp in the Cw*0428 allele.

[image: image1.emf]Well No.

1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24

HLA-Cw allele

1

*04010101-040105

1 2 3 11

*0403

1 4 11

*040401-040402

2 3 f 11

*0405

1 2 3 6 11

*0406

4 f 11

*0407

1 2 7 11

*0408

1 2 3 w 11

*0409N

1 2 3 9 11

*0410

1 3 10 11

*0411

1 3 10 12

*0412

1 w 3 11 13

*0413

2 3 f 11 14

*0414

2 3 11 15

*041501

1 2 3 11 16

*041502

1 2 3 11 20

*0416

1 11 14

*0417

1 2 3 11 17 20

*0418

1 2 3 11 14

*0419

w 2 3 11 18

*0420

1 2 3 11 19



[image: image2.emf]*0423

1 2 3 11 21

*0424

1 2 3 11 22

*0425

1 2 3 11 23

*0426

1 2 3 11 24

*0427

1 2 7

*0428

1 2 3 11 15

*0429

1 2 3 11 12

*0430

1 2 3 9 11 12

*0431

1 2 3 10 11

*0432

1 2 3 7 11

*010201-0120, 020201-020203, 020205-

020206, 0204-0211, 0213-0215, 0217, 

0219-0221, 050101-050104, 0503-0510, 

0512-0516, 0518, 0519, 06020101-0603, 

0606-0617, 0802, 0805, 0807, 0812, 

0815, 120201-1213, 1215-1217, 1221, 

140201-1404, 1407N, 1408, 1508, 16040

1

*0203, 0216, 0218, 0511, 0517, 0604, 

080101-080102, 0803, 0804, 0806, 0808-

0811, 0813, 0814, 0816, 1214, 1218, 

1220, 1406, 150201-1507, 1509-1513, 

1516-1521, 1701-1704

5

*0212

1 4

*0305, 0325, 0327

16

*0334

7

*0605

1 11

*0708

14

*1409

1 16

*1515

5 8

B*5802

5 8

Length of spec. PCR product(s)

2

250 220 145 210 250

95

145 270 180 195 180 125 225 155 170 130

90

225 155 540

85

130

85

170

220 220 180 190 190

Well No. 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24

Specific Reaction

1st match: Cw*

2nd match: Cw*

Control Reaction

Failed Controls

False Positive

False Negative


1The nucleotide sequence of the Cw*0402 allele has been shown to be identical to Cw*04010101.

The nucleotide sequence of the Cw*0421 allele has been renamed to Cw*041502.

The nucleotide sequence of the Cw*0422 allele has been shown to be identical to Cw*0421.

1Primer mix 9: Specific PCR fragment of 180 bp in the Cw*0430 allele. Specific PCR fragment of 220 bp in the Cw*0409N allele.

Primer mix 10: Specific PCR fragmentof 195 bp in the Cw*0431 allele. Specific PCR fragment of 220 bp in the Cw*0410 and Cw*0411 alleles.

Primer mix 12: Specific PCR fragment of 125 bp in the Cw*0411 and Cw*0429 alleles. Specific PCR fragment of 180 bp in the Cw*0430 allele.

Primer mix 14: Specific PCR fragment of 155 bp in the Cw*0416 allele. Specific PCR fragment of 190 bp in the Cw*0413, Cw*0418 and Cw*0708 alleles.

Primer mix 15: Specific PCR fragment of 170 bp in the Cw*0414 allele. Specific PCR fragment of 190 bp in the Cw*0428 allele.

‘w’, may be weakly amplified.

‘f’, may be faintly amplified.

Changes in revision R01 compared to R00:

1. The Cw*040101-040102, Cw*0406 and Cw*0413 alleles are faintly amplified by primer mix 5.
Changes in revision R02 compared to R01:

1. The Cw*0408 allele is weakly amplified by primer mix 8.

2. The numbering of the foot notes in the specificity table has been corrected.







Olerup SSP AB

www.olerup.com
Hasselstigen 1

SE-133 33 Saltsjöbaden

Sweden

