# Gel Documentation Form and Worksheet

# KIR HLA Ligand (104.201-12/12u)	                            Lot: 7L0                                            Expiry Date:   2024-10-01

# Name:____________________

# Sample ID:________________

# DNA Conc.(ng/ul):___________

# Test  Date: _______________

# Tested By: _______________

# Review Date:_______________

# Reviewed By:_______________

# Interpretation:___________	Failed lanes: _____________ 	 Comments:_________________

Gel Picture

## PHOTO DOCUMENT

Abbreviations

ICB: Internal Control Band

AmpS: Amplicon size

Notes: Product sizes are approximate. For detailed information, see the lot-specific Specificity Table and Interpretation Table.

This table is intended as a guide. For interpretation always use the Interpretation Table and/or Specificity Table.

A minor subset of rare B and C alleles with the listed target motifs will not be amplified by the kit due to downstream sequence variations.

Primer mixes 3 and 4 may have tendencies of unspecific amplifications.

Primer mix 8 contains a negative control, which will amplify a majority of HLA amplicons as well as the amplicons generated by the control primer pairs matching the human growth hormone gene. HLA-specific PCR product sizes range from 75 to 200 base pairs and the PCR product generated by the HGH positive control primer pair is 200 base pairs.

This lot will not amplify the C*01:46 allele in primer mix 1.

The alleles C*05:32 (80Arg), C*15:60 (80Gln) and C*15:71(80Ile) will be amplified in primer mix 2.

1HLA-A, HLA-B and HLA-C alleles listed on the IMGT/HLA web page 2020-July-13, release 3.41.0, www.ebi.ac.uk/imgt/hla.

2Alleles that have been deleted from or renamed in the official WHO HLA Nomenclature up to and including the last IMGT/HLA database release can be retrieved from web page http://hla.alleles.org/alleles/deleted.html.

Abbreviations

w: might be weakly amplified.